## Slide 1: Alignment, Scanning
Malcolm Guthrie
www.europeanspallationsource.se
13 September, 2018

## Slide 2: Alignment needs
Traditional alignment of neutron samples done with a tape measure
Majority of ESS instruments consider using < 1mm samples
Need higher precision alignment approach.
Notes: SAM = survey, alignment & metrology

## Slide 3: synchrotrons
Luckily, problem solved in 1980’s
Scanning based approach.
Notes: SAM = survey, alignment & metrology

## Slide 4: Scanning methodology
x
x
x-axis
Beam
Detector
z-axis
Setup:
Notes: SAM = survey, alignment & metrology

## Slide 5: Requirements
Monitors
Active area larger than the beam
Uniform response across area
TOF capable
Motors
High speed: full scan in < 10s
2-axis
Complex scans (e.g. spirals)
Notes: SAM = survey, alignment & metrology

## Slide 6: Requirements
Software
Easy user interface
Relative vs absolute scan capability
ability to define scan parameters (start, stop, step size)
Interfaces
SAD
Detector group
ICS group
Motion control…
Notes: SAM = survey, alignment & metrology

## Slide 7: Current status
Prototype in Utgård, (using laser)
Diode based detector
Linear motor (driven by Beckhoff controller)
Next steps:
Scanning software (already exists in EPICS)
Neutron version (V20)
Notes: SAM = survey, alignment & metrology

## Slide 8: Beam characterisation
In addition to alignment, also reveals beam profile
TOF capability enables wavelength-dependent measurements
Scanning at different points along beam measures divergence + focal points
Notes: SAM = survey, alignment & metrology

## Slide 9: Alignment of equipment
neutron pinhole/cross-hair is also optically visible
Can align a camera system to same target, thus mapping optical camera to neutron beam
Notes: SAM = survey, alignment & metrology

## Slide 10: Metrology
Need high-precision metrology system to commission mounts.
SAM provides alignment services (laser-based)
Need local system for day-to-day commissioning
Devised “simple” 2-camera system:
Cam 1
Cam 2
O
Beam
Cam 1
z
y
Cam 2
z
x
after calibration: click gives x,y,z coords in USCS
Notes: SAM = survey, alignment & metrology

## Slide 11: “Simple”
Devise optical spec for cameras (F.L., FOV, mag)
Devise interface spec for cameras (PoE)
Implement EPICS interface for cameras
Software to convert pixel coords to USCS coords.

## Slide 12: Camera test

## Slide 13: Camera test
Coordinate | Original pos. (mm) | Shifted pos (mm) | δy laser(mm) | δy camera (mm)
X | 0.24 | 0.47 | + 0.23 | ?*
Y | -0.33 | 0.06 | - 0.39 | -0.390
z | -0.01 | 0.01 | -0.02 | -0.027
< 10μm precision on displacements

## Slide 14: Universal Coordinates
Kinematic Positioning
Interface has two separable parts:
High positional repeatability when they are connected:
High Level Requirements:
Rapid installation, minimise online alignment
High positioning precision ~1mm sample sizes).
Flange and floor-mounted compatibility.
Minimise adapters by establishing standard interfaces
Concepts:
Use kinematic positioning to ensure repeatability
Calibrate instrument part of mount to have identical spatial relation to the sample position on each instrument
Repeat process on multiple ”Levels” to cover wide range of equipment and tolerance requirements
Equipment mount
Instrum. mount
Equipment mount
Instrum. mount
Calibration process:
Take a unique calibration standard and interface it with each instrument
Translate/rotate instrument mount until standard is centred on sample position
Lock this position

## Slide 15: Alignment protocol* (*assumes known sample position “+”)
Inst. 1
Inst. 2
Inst. “n”
Before alignment:
Beam
Calib’Standard

## Slide 16: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Beam
Calib’Standard

## Slide 17: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Beam
Calib’Standard

## Slide 18: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Beam
Calib’Standard

## Slide 19: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Beam
Calib’Standard

## Slide 20: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Etc…
Beam
Calib’Standard

## Slide 21: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Etc…
Beam
Calib’Standard

## Slide 22: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Etc…
Beam
Calib’Standard
Expect periodic recalibration (beam drift, concrete subsidence etc.)
Note: Samples must be calibrated relative to “equipment” mount
This can be done offline by calibrating an “alignment station” in exactly the same way as the instruments

## Slide 23: Hidden ROI
Define sample frame As
  A Cartesian frame, defined by three basis vectors, fixed with respect to sample
  As = [e1,e2,e3]
  Any point in the sample described by vector r, written in terms of As as
  X
  Where X = (x,y,z) is the coordinate matrix
Coordinates of ROI are known in As
Also coordinates of several “fiducial” points* are known

## Slide 24: Hidden ROI
2. Equivalent instrument frame AI
  With coordinate matrices X’ = (x’,y’,z’)
3. Determine X’ for several (visible) reference points on object.
4. Obtain transformation between AS and AI
5. Obtain coordinates of ROI in instrument frame
AI
Sample position (Origin of instrument frame)