## Slide 1
MEBT Controls
ESS-Bilbao
Control & Beam Instrumentation Group
Idoia Mazkiaran Zelaia
imazkiaran@essbilbao.org
November – 2018

## Slide 2
2
Summary
Introduction
Faraday Cup: FC
Emittance Meter Unit: EMU
Scrapers
Buncher
Fast Chopper

## Slide 3
ESS MEBT
3

## Slide 4
4
MEBT General Layout

## Slide 5
Faraday Cup HW: Control schematic

## Slide 6
Faraday Cup HW: µTCA platform
Main control IOC and signal acquisition.
µTCA crate: Schroff custom 3U crate up to 6 AMC cards, 4 RTMS, redundant Power Supplies, 1 MCH slot.
MCH: NAT-MCH-PHYS-R1
CPUs:
  IOxOS IFC_1410 QorIQ T2081 & Kintex UltraScale AMC.
  Concurrent Technologies AM 900/412-99.
  Digitizer board:
    IOxOS FMC ADC 3111, 8 analog channels, 250 MSPS, 16 bits.
  Event Receiver board: MRF MTCA EVR-300U.
External integrated systems controlled from
the previous crate HW:
  FC insert/retraction: solenoid valve.
  Repeller control.
  Interlocks interface.
Siemens PLC for cooling and slow interlocks functions.

## Slide 7
Faraday Cup: Front End / Back End
Front End SPECIFICATIONS
Number of channels 1
Nominal input current 62.5 mA
Maximum input current 100 mA
Minimum input current 6.25 mA
Bandwidth > 2MHz
Slow tuning mode 50us at 1 Hz
Fast tuning mode 5 us at 14 Hz
Output: Differential +/- 10 V, High Impedance
ADC Resolution 16 bits
Back End:
Auxiliar Power supplies for Front End.
FC Repeller.

## Slide 8
8
SW general control
FC SW based in EPICS: E3 Environment.
Four IOCs running.
Main FC IOC mebt_fc module:
General control and flow of the FC, Repeller control.
Modules:
  Ifc14edrv, ifc1410fc: Digitizer driver modules.
  Other modules:
  asyn, nds3, nds3epics.
2. Synchronization IOC: EVR received event generates µTCA back plane
acquisition trigger signal.
3. Ethercat IOC:
  ecat2 module: Beckhoff ethercat modules control.
4. Slow Interlocks IOC:
  sp7plc module: PLC interconnection module.
  MEBT_plc_FC: Interlocks specific developed module.
Notes: 8

## Slide 9
9
FC Intermediate Control
1. FC Repeller. (Back-end)
ISEG DPR 10 106 24 10 ESH
Vout: 0 to +/–1000 V
Imax: 10 mA
Vin/Iin: 24 V/0.8 A
Control Signal 0-10 V
2. FC insert/retraction: pneumatic valve.
Model: Solenoid valve 5/2 Monostable - FESTO – VUVG-L18-M52-MT-G14-U-1K7L
Home and end run switches: Local Switches: Proximity Sensor - Festo - SME-8-K-24-S6 (161756)
MPS switch: MPS Switches: Crouzet - 83 133 035
3. Auxiliar Powers Supplies. (Back-end) Front End Power Supplying.
4. Health Status, Beam Permit.
Notes: 9

## Slide 10
10
FC Engineering GUI
Notes: 10

## Slide 11
Faraday Cup LPS (Local Protection System)
  .
Siemens PLC for cooling and other slow interlocks functions.
Faraday Health: General
Repeller status.
Cooling: PT100.
Water flow sensor.
PLC 24 Vdc distribution.

## Slide 12
12
FC test stand at ESS Bilbao ISHP
Configuration | (A) | (B)
Power (W) | 685 | 690
H2(%) | 45 | -
t (μs) | 1600 | -
Rep rate (Hz) | 1 | -
Plasma PS (A) | (2;5;10’2;3’5) | (2;5;10’2;3’5)
Stud | (575;632) | (519;636)
Absorbed power (%) | 98’67 | 99’1
Solenoid (A) | 150 | -
HV (kV) | 45 | -
Repeller (kV) | -1’5 | -
Ion Source HP configuration
Front End: 100 Ohm resistor.
Front End: FC active electronic FE.
NI Synchronization
Collimator used to protect FC
Publication in process

## Slide 13
Slit & Grid system schematics
13

## Slide 14
EMU HW: µTCA platform
µTCA platform for EMU prepared for E3 EPICS environment.
Main control IOC and signal acquisition.
µTCA crate: Schroff custom 3U crate up to 6 AMC cards, 4 RTMS, redundant Power Supplies, 1 MCH slot.
MCH: NAT-MCH-PHYS-R1
CPUs:
  IOxOS IFC_1410 QorIQ T2081 & Kintex UltraScale AMC.
  AM 900/412-99 by Concurrent Technologies.
  Digitizer boards:
    2 IOxOS FMC ADC 3117, 20 analog channels, 5 MSPS, 16 bits.
  Event Receiver board: MRF EVR-300U.
External integrated systems controlled from the previous HW crate :
  Motion Control: Bus ethercat, linear motion, Beckhoff modules.
  Power supply control for BIAS voltage supplying.
  Additional control functions: plane change, power supplying checking, wire checking, calibration.
  Analogue front-end interrelation for signal conditioning.
  Interlocks interface.
Slow Interlocks Siemens PLC for cooling and slow interlocks functions.
( Developed in EEE over VME, some needs to migrate to E3)
14

## Slide 15
EMU: Front End / Back End
Front End:
Number of channels 24
Minimum input current 100 (200) nA
Maximum input current 100 (200) µA
Bandwidth ≥ 2MHz
Slow tuning mode 50us at 1 Hz
Fast tuning mode 5 us at 14 Hz
Output: Differential +/- 10 V, High Impedance
ADC Resolution 16 bits
Modular Sub-rack with backplane connections
  + Large modularity.
  + Robustness (no internal cabling)
  - Complex assembly design
Back End:
Auxiliar Power supplies for Front End.
BIAS Power Supply.
Features:
1ST STAGE: Transimpedance
2nd STAGE: Voltage Amplification and single-ended to differential
3rd STAGE: Filter and Transmission line driver
Slot 1: Outputs.
Slot 2: Tests control.
Slot 3: Power.
Slot 4: Inputs: EMU Plane.
Slot 5 .. 8: Signal accommodation channels (4 boards x 6 channels).

## Slide 16
16
EMU SW development
Main systems in the EMU are controlled by a μTCA chassis:
System SW flow lead by sscan record.
Main Features:
Signal acquisition: 2 planes with 24 channels each sharing same digitizing HW.
Equidistant positions:
  Precalculated positions. Positions stored in file.
Motion Control: 4 motors: 2 slit blades, 2 grids (24 tungsten wires).
Grid recursive movement over slit movement.
Linear motion and incremental encoders, two limit switches ( two for MPS).
Motion axis from for slit and grid of each plane correlated.
      Stepper motor Nanotec ST4118M1804-B. Brake Nanotec BWA-025-5.
Incremental encoder SIKO MSK1000-11-M-E1-2-I-1-1.00 : resolution 1 µm.
Movement Based on sscan record:
Plane changing.
Wire Checking: 10 Vdc onto tungsten wires.
States:
Before slit movement.
Before grid movement.
On Grid scanning.
After all points of Grid scanned for an specific slit position.
After slit scanning: Emittance calculations.

## Slide 17
17
EMU SW development (II)
Information arrays obtained:
  x, y: slit position array.
  x’, y’: angle of grid wires position respect slit position array.
  Avg: average in volts of the current induced by the beam in each wire.
  Emittance RMS.
  Twiss parameters: alpha, beta and gamma.
EMU SW developments based in EPICS: E3 Environment.
Four IOCs running.
Main FC IOC mebt_emu module: sscan control and functionalities.
Digitizer driver modules, other modules: asyn, nds3, nds3epics.
2. Synchronization IOC: EVR Event generates µTCA back plane acquisition trigger signal.
3. Ethercat IOC: Based on ecmc module: Analog and digital I/O, motors control.
4. Slow Interlocks IOC:
  sp7plc module: PLC interconnection module.
  MEBT_plc_EMU: Interlocks specific developed module.
Notes: 17

## Slide 18
18
EMU Engineering GUI
Notes: 18

## Slide 19
19
EMU LPS
Notes: 19

## Slide 20
20
Scraper Control Schematic

## Slide 21
21
Scrapers motion features
6 Scraper axis to set in horizontal and vertical pairs.
Movement EPICS development either as independent axis or as correlated axis by pairs based on ecmc module.
Actuator: Stroke: 80 mm.
Pitch: 2 mm.
Total required torque in motor axis: 0,42193 Nm.
Stepper motor Nanotec ST4118M1804-B.
Brake Nanotec BWA-025-5.
Incremental encoder model SIKO MSK1000-11-M-E1-2-I-1-1.00 with magnetic band SIKO MSK1000-1-0.3-10-0.01-St-TM-AM-0: resolution: 1 µm, index signal 1 mm.
Bekchoff modules for motion control.
Motor features:
Motor features:

## Slide 22
Scrapers limits and homing
Actuator limits
Mechanical limits: A, B.
Limit switches: F: Initial SC Travel Limit Switch.
D: MPS Limit Switch.
C: Final SC Travel Limit Switch.
SW limits.
Homing procedure: Home: home limit switch (or encoder reference signal).
Switch response at 1 mm/s
 | Zero to ONEs Statistics (µm) | 
 | 10 mm/s | 1 mm/s
Average | 239 | 160
Max. Error | 3 | 33
Std. Deviation | 3 | 18
Jitter | 6 | 53
Home position accuracy:
Based only in limit switch: > 50 µm.
Based on index pulse of encoder each 1 mm: 1 µm.

## Slide 23
23
Scraper Motion Engineering GUI
EPICS developments based on ecmc module.
Beckhoff modules:
EL1018: Digital Inputs: limit switches, MPS home switch and MPS Beam Permit signal.
EL2808: Digital Outputs for powering switches to 24V, motor brake and Health Status signal to MPS.
EL5101: Encoder driver.
EL7047: Motor driver: 2000 Hz.
Max speed: 20 mm/s.
Max required speed: 10 mm/s.

## Slide 24
24
Buncher Tuner Mobile
3 Mobile tuners for each Buncher, used at cavity tunning.
Movement EPICS development either as independent axis or as correlated axis by pairs based on ecmc module.
Actuator: Linear shift mechanism: LSM64-21395-001 (UHV).
Stroke:75 mm.
Pitch: 2,54 mm.
Torque to move 245N load about 0.75Nm.
Stepper motor McLennan 23HT18C230.
12.5:1 spur gearbox.
Renishaw LM10IC010EA10A00 magnetic Encoder: 10µm resolution, index signal 2 mm.
Bekchoff modules for motion control.
Limit Switches Home Sensor.

## Slide 25
25
Buncher Tuner Mobile
EPICS developments based on ecmc module. Additional BU_motion_interface module.
Beckhoff HW modules:
EL1018: Digital Inputs: limit switches, home sensor.
EL2808: Digital Outputs for powering switches to 24V.
EL5101: Encoder driver.
EL7047: Motor driver: 8000 Hz.
Max speed: 5 mm/s.
Max required speed: 2,5 mm/s.

## Slide 26
26
Fast Chopper schematics

## Slide 27
27
Fast Chopper Timing
Normal operation chopper signals
Edge of trigger signals
Triggers to introduce in the pulsers
Output pulses from the pulsers
Trigger Width: Recommended triggers from 0,1 to 1 s.
distance between triggers from 5 s to 3 ms.
Offset Delay: Synchronize chopper deflection with beam (µs).
Fine Delay: From 2,2 ns to 12,4 ns in steps of 10 ps.
Distance between paired triggers: From 5 s to 3 ms.
Rising edge of the trigger signal that is a few ns.

## Slide 28
28
Fast Chopper SW
Control developments in EPICS based on the EEE’s recommendation.
Drivers needed for HW and VME :
IOxOS VME drivers
MRF (Timing) drivers.
Those are the EPICS modules used in the SW solution:
Timing control:
EVGChopper module: event generator configuration.
EVRChopper module, for configuring the event receiver of the chopper.
mrfioc2: MRF timing module special version for UNIV-TTL-DLY support.
System and device support related modules: pev, pevdrv, devlib2, regdev.
Pulser control:
PulserProtocol module, for configuring the features of the pulsers. (IOC for each pulser).
Serial communication and related modules: asyn, streamdevice, calc, seq, pcre.
Slow Interlocks:
Chopper_plc module, used s7plc module.

## Slide 29
29
Fast Chopper Engineering GUIs
Chopper control of the pulsers (FID GmbH) through proprietary protocol.

## Slide 30
30
Thank you for your attention.
Tack så mycket!
¡Muchas gracias!
Eskerrik asko!