Description:	This document describes the design and current state of software and FPGA firmware for the ionisation chamber-based Beam Loss Monitoring system (icBLM) of the ESS accelerator
Title:	Software and Firmware Design Report for icBLM

Authors:	Grzegorz Jabłoński
Co Authors:	Wojciech Jałmużna, Rafał Kiełbik, Wojciech Cichalewski, Andrzej Napieralski
Version:	1.0
Date:	2019.02.05

Revision Table

Revision	Date	Modification
1.0	2019.02.05	Final version provided for review
1.1
1.2
1.3
1.4
1.5
1.6
1.7
1.8
1.9
2.0

Table of Contents

# Introduction
The primary goal of the Beam Loss Monitoring (BLM) system is to monitor beam losses of the ESS accelerator and detect abnormal beam behaviour and promptly inhibit beam production in case of beam failures to keep the machine safe from beam-induced damage. icBLM system is the main BLM system in the superconducting parts of the linac and is based on ionisation chambers. Its specification can be found in [1].
The works summarised therein has been performed in the frame of Polish in-kind contribution to ESS [2].
The Workunit 2 (icBLM and nBLM controls and realization) comprises:
-	Development of FPGA firmware that runs on the icBLM and nBLM back-end electronics (BEE) and has the following tasks:
o	Process the incoming data (from ADC or other front-end electronics)
o	Provide the data to the (EPICS) control system for monitoring purposes
o	Run algorithms for machine protection purposes that assess the beam conditions and in case of unacceptable beam losses inform the Beam Interlock System by dropping the BEAM_PERMIT signal and also providing data that is needed to analyse the conditions that led to dropping the beam permit
o	Perform checks to verify that the system components and functions are working properly.
-	Development and implementation the following software:
o	icBLM specific algorithms for data processing that run on the icBLM BEE CPU,
o	The low-level drivers to interface to the (EPICS) control system,
o	Implementation of EPICS databases to store and process the data,
o	Development of graphical user interfaces to visualize the icBLM specific data.
Software development for the nBLM system is covered by a separate contract and is not in the scope of this contract. Nevertheless, the software demonstrating the functionality of the nBLM software and allowing its performance evaluation has been also provided.
# Current implementation status
## Hardware and software platform
The hardware platform consists of the IFC_1410 FMC carrier AMC, equipped with Xilinx Kintex UltraScale KU040 FPGA and PICO4 - FMC extension card used with IFC1410 equipped with 4 1 MHz isolated channels for micro-current sampling. The firmware has been implemented within the TOSCA framework, provided by the board manufacturer, IOxOS. There are two option for the controlling CPU – the POWER processor on the AMC module or the Concurrent AM90x/41x in the MTCA crate.
## Currently implemented firmware blocks
The implemented firmware blocks include:
- Data acquisition from ADC (4 channels)
- Buffering of data-on-demand
## Firmware blocks not implemented
The firmware blocks that are not implemented include:
- Preprocessing
- Baseline calculation
- Modulation detection
- Protection function block
- FBIS interface block
- Smart trigger block
- Low Latency Link block
- Backplane triggering
## FPGA resource utilization
Resource	Utilization	Available	Utilization %
LUT	49577	242400	20.45
LUTRAM	3556	112800	3.15
FF	61373	484800	12.66
BRAM	197.50	600	32.92
IO	337	520	64.81
GT	8	20	40.00
BUFG	22	480	4.58
MMCM	3	10	30.00
PCIe	2	3	66.67

Due to relatively low requirements on speed of the algorithm/acquisition and low numerical complexity of required algorithms, no resource problems and/or bottlenecks are expected.
## Software implementation status
Currently only a command line application capable of raw data acquisition is available.

# Past problem and issues
## DDR Calibration
The TOSCA framwork has had a bug, causing segmentation violation errors on the host CPU during the on-board DDR memory calibration procedure most of the time. Several reboots of the IFC_1410 modules were needed to make the board operate properly. The problems became more severe after migration of the Vivado version from 2016.4 to 2018.2. The problem has been finally identified by IOxOS and fixed in December 2018.
## Timing closure problems
The examples provided by IOxOS did not synthesize without timing errors “Out of a box”. Most implementation runs resulted in pulse width timing violations related to the 250 MHz clock driving the DDR memory controller. Finally we have been able to fix the problem by constraining the global buffer output driving this clock to a specific location.
## DMA performance problems
Initially we were instructed to use the Tosca kernel driver for the communication between CPU and FPGA. This driver provided very low DMA throughput of order of 550 MiB/s. Later on we have transitioned to the ESS-Tsc driver with the throughput of 950 MiB/s.
# Current issues
## Loss of data during simultaneous DMA transfers from both DDR banks
There is a bug in the firmware or device driver causing data loss when the DMA transfer is performed simultaneously from two different memory banks using different DMA channels from different threads. The workaround is protections of the DMA transfer function in the userspace program by mutex, but this solution reduces performance.
## Tsc driver not supporting scatter-gather operations and prone to resource leaks
Tsc driver has high performance, but requires allocation of contiguous physical memory buffers for operation. When memory is fragmented, it is not possible to allocate these memory buffers and a CPU reboot is required. Tosca driver did not have this problem, but it offered much lower performance.
These memory buffers have to be manually allocated and are not freed automatically when the device file is closed. It requires appropriate signal handlers in the userspace program to avoid resource leaks.
## Problems with PCIe Gen3 link to Concurrent CPU
The PCIe link speed between Concurrent CPU and the IOxOS board alternates between Gen1 and Gen3 after each system reboot.
## No separate interrupt support for both memory banks
Only one interrupt, common for both memory banks, is supported. It is a firmware issue.
## Tsc Driver cannot access PON space when run on Concurrent CPU
TscMon ( and custom software applications ) fail to access registers placed on PON configuration space ( such as Power On for FMC cards ) when run on Concurrent CPU. The same software and TscMon commands can be successfully run on embedded IFC1410 processor. The bug prevents proper behaviour of the software when only one PCIe endpoint (to Concurrent) is present in the design.
# Future plans
As a first step, the preprocessing, baseline calculation and modulation detection blocks will be implemented. Next, the EPICS IOC and operator panels will be designed.
# References
[1]	Requirements and Technical Specifications - icBLM System (https://jira.esss.lu.se/browse/BIG-1246)
[2]	Schedule technical annex to the in-kind contribution agreement signed between European Spallation Source and Department of Microelectronics and Computer Science, Lodz University of Technology on 2016-11-08; CHESS id: ESS-0149090