## Slide 1: icBLM verification plan
Clement Derrez
European Spallation Source ERIC
Date

## Slide 2: Reminders: Naming convention, data management tools
Naming convention: Every Field Replaceable Unit (FRU) is named. A FRU can be a simple part or an assembly
Data management system: going from Insight to EAM
  Data export / import with other software tools
  Ensures traceability between tests and production data,
  system components and laboratory devices.
  Systems and subsystems tests results coming from
  different locations (ESS, IK Partner, commercial partners…)
  SciCat is a good candidate to complete EAM
Data files format: HDF5
  Files, groups and datasets minimum mandatory
  attributes are defined
Insight -> EAM
Component requirements
System FRUs list
Tests report
Test plan
Laboratory devices
Test Design document
JIRA: timeline

## Slide 3: Reminders: acceptance workflow
System requirements
System Verification plan
Validation of verification plan
System requirements are analyzed for each system and its components
Verification document is validated by system lead: we make sure test requirements are correctly translated
Test coverage is reviewed.
Describes subsystems' test setup and characterization procedures
Tests IDs defined for each measurements
When not COTS: test plan is prepared/discussed and agreed with IK partner
Run tests
Acceptance tests performed in BI lab, possibly compared with IK results.
System verification performed after installation, without beam and with beam from the LCR
System hand over

## Slide 4: System FRUs
Component ID | Acronym | Description
AMC | AMC | 1 per up to 2 FMC, IOxOS 1411 type
FMC Digitizer | FMC | CAEN FMC-PICO-1M4-C3
Timing receiver | EVR | 1 per MTCA chassis, event receiver for trigger, clock, calibration announcement and beam/machine mode information
Chassis | MTCA | 1 per system
MicroTCA Carrier Hub | MCH | 1 per MTCA chassis
MicroTCA Power Supply | PS | 1 per MTCA chassis, 600/1000 W
MicroTCA CPU | CPU | 1 per MTCA chassis, intel i7
ECAT crate, populated | ECAT | 1 per HV supply
HV supply | HV | 1 per n systems, ISEG HV supply
Rack patch panel | PP | 1 per n systems, top of rack
Rack | U | 1 per n ACCTs, rack for electronics
Detector unit | icBLM | 1 ionization chamber
Beam line patch panel | BLPP | 1 per N ionization chambers

## Slide 5: Verification identification
Test ID | Verification procedure
icBLM_1 | Visual inspection. Check visually the HV card delivered
icBLM_2 | Install HV unit with ECAT. Connect to IC with test cable. Apply test voltage. Verify communication between ECAT and HV Supply
icBLM_3 | Verify Current readback
icBLM_4 | Verify Voltage readback
icBLM_5 | Verify Voltage readback RMS
HV unit and ionization chamber detectors
Test ID | Verification procedure
icBLM_6 | Visual inspection. Check visually the Detector: SHV connectors, BNC connector, Serial number labels.
icBLM_7 | Connect Signal cable to pico-ammeter in leakage testbench. Run automated measurement
icBLM_8 | Connect Signal cable to pico-ammeter in Rad Source testbench. Run automated measurement
HV Unit
Detectors

## Slide 6: Verification identification
HV unit and ionization chamber detectors
Digitizer
Test ID | Verification procedure | Threshold
 | Run auto test software to measure: | 
icBLM_9 | Channels RMS noise, OC, RNG0 | < 250 nA
icBLM_10 | Channels RMS noise, OC, RNG1 | < 15 na
icBLM_11 | Channels RMS noise, with long cable and detector, RNG0 | < 150 nA
icBLM_12 | Channels RMS noise, with long cable and detector, RNG1 | < 180 nA
icBLM_13 | Bandwidth | 300 KHz
icBLM_14 | ADC ENOB | 14.5 bits
icBLM_15 | Crosstalk | < 0.1% FS

## Slide 7: Reminders: After-installation tests sequence
After-installation tests sequence:
  cold check-out
  commissioning with beam
  quick self-check
Sequential testing is important (tests timestamps are checked automatically). Otherwise the statement: “the instrument is installed and working properly” has less confidence.
During debugging:
  Relevant tests in each architectural layer are repeated until satisfactory results are obtained.
  If the problem is identified and can be isolated within its layer, there is normally no need to repeat all of the tests which are sequentially following. Depending on the situation, some test might become mandatory nevertheless (for instance a software recompilation might entail a standard interface check, and a repaired connector might entail a signal transmission check).
Notes: Cold check out: Testing of the whole instrument on all architectural layers:
  monitor, front-end electronics, cables, mTCA-electronics, timing, data treatment, publishing, network transmission and machine protection interface.
Commissioning with beam: Aims at verifying the correctness of the integration for machine operations.
  Includes initial comparisons and cross-calibrations in order to gain confidence in the instrument. Performance limiting factors are identified.
Quick self test: The self-test procedure includes testing of calibration, machine protection and data transmission.

## Slide 8: Verification identification
Cold check out
Assembled system: detector, uTCA electronics, ECAT electronics Cold check out Part ID: NA |  |  |  | 
Detailed tests results document: links on CHESS, if applicable. we use the control room OPI for all these steps from now on |  |  |  | 
Test ID | Verification procedure | Threshold | Result | Pass / Fail
icBLM_16 | Check signal RMS and pkpk noise in RNG1 | < 180 nA |  | 
icBLM_17 | Self -test verification: Check signal continuity through the HV modulation signal |  |  | 
icBLM_18 | Detector mapping and complete acquisition chain verification: Use a plastic hammer to induce a shock to the detector. Verify the corresponding signal peak on the OPI |  |  | 
Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  | 

## Slide 9: Verification identification
Commissioning with beam
Commissioning with beam Part ID: NA | 
Detailed tests results document: links on CHESS, if applicable. The test requires RF (with stable settings - not necessary nominal) present. | 
Test ID | Verification procedure
icBLM_20 | Trigger delays / Verify background subtraction using offline analysis procedure
icBLM_21 | Trigger delays / Monitor accumulated loss (or neutron counts) over the beam pulse: For each trigger delay setting an average value for this loss is computed Report the trigger setting with highest average value
icBLM_22 | Monte Carlo Simulation verification and equivalent loss scaling factor: Define a set of controlled losses for simulation verification Compare simulated and measured results. In case of larger discrepancies find the source and modify simulation model accordingly. Once the simulation geometry model is verified, use the results to define scaling factors. For each detector connect number of lost protons to measured current for particular loss scenario. This gives a factor that can be used to calculate "equivalent lost protons" from the measurement during operation. Here each group of detectors has a loss scenario assigned to.
icBLM_23 | Protection function commissioning Identify controlled loss scenarios with different loss time evolution (different time constants) and tune the protection function algorithm Produce these controlled losses to tune the protection algorithms
icBLM_24 | MP thresholds Identify a few likely accidental scenarios that are most damaging Use Monte Carlo simulations (Geant4) coupled with thermo-mechanical simulations (ANSYS) to understand damage potential. Produce risk matrix which serves as a bassline to select MP thresholds.
Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) | 
Notes: Trigger delays / Verify background subtraction, offline analysis procedure:
For each trigger setting: raw data without baseline subtraction (one of the Data on Demand-DoD buffers) is acquired. Average baseline is subtracted (calculated from another DoD buffer where waveforms that are used for baseline subtraction are stored).
For each pulse period (14Hz) average baseline waveform from the raw data is subtracted to obtain "processed" waveform.
From at least 100 "processed waveforms" an average is computed. Report average processed waveform with the best trigger setting.
Trigger delays / Monitor accumulated loss (or neutron counts) over the beam pulse: (ie. loss over BEAM_ON period inside the pulse period):
For each trigger delay setting an average value for this loss is computed
Report the trigger setting with highest average value
Monte Carlo Simulation verification and equivalent loss scaling factor:
Define a set of controlled losses for simulation verification
Compare simulated and measured results. In case of larger discrepancies find the source and modify simulation model accordingly.
Once the simulation geometry model is verified, use the results to define scaling factors.
For each detector connect number of lost protons to measured current for particular loss scenario. This gives a factor that can be used to calculate "equivalent lost protons" from the measurement during operation.
Here each group of detectors has a loss scenario assigned to.
Protection function commissioning
Identify controlled loss scenarios with different loss time evolution (different time constants) and tune the protection function algorithm
Produce these controlled losses to tune the protection algorithms
MP thresholds
Identify a few likely accidental scenarios that are most damaging
Use Monte Carlo simulations (Geant4) coupled with thermo-mechanical simulations (ANSYS) to understand damage potential.
Produce risk matrix which serves as a bassline to select MP thresholds.

## Slide 10: Coming priorities
Continue with tests in FREIA to better assess cables’ noise performances and background levels
Further integration tests with the ISEG HV PSU in ICS integration lab
Deeper evaluation of the picoammeter FMC on the IOxOS carrier
Detector tests in BD lab

## Slide 11
Thank you!
Questions ?