## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development
Status of the FREIA Test Stand
Roger Ruber for the FREIA Team
aTAC, ESS, 11 April 2019

## Slide 2: FREIA Laboratory
cryogenics
- liquid helium
- liquid nitrogen
control room
- equipment controls
- data acquisition
radio-frequency (RF) power sources
3 bunkers
with test stands
horizontal cryostat
vertical cryostat
State-of-the-art Equipment
Facility for Research Instrumentation and Accelerator Development
Competent and motivated staff
collaboration of physics (IFA) and engineering (Teknikum).

## Slide 3: Double Spoke Cavity Package (2017)
Test from March – May 2017
Warm
  RF conditioning with IPNO then FREIA system
    about 30h; multipacting in 20-60 kW region
    max 120 kW forward power (limit set by IPNO)
Cold
  RF conditioning with FREIA system
    about 30h; 3 major multipacting regions
  RF testing
    Q0 = 2.6 x108 at 9 MV/m (w/ heat load correction)
    Lorenz force detuning ~400 Hz at 9 MV/m
    cavity tuning sensitivity = 150 kHz/mm@2K
    pressure sensitivity = +27.1 Hz/mbar
Final status/results
  results presented at SLHiPP-7 (9 June 2017)
  full report published: urn:nbn:se:uu:diva-335434

## Slide 4: High-β Elliptical Cavity Package (2018)
1st run in June 2018
Warm-up and open to fix tuner problem
2nd run in August 2018
  cooldown, then RF re-conditioning at cold
  measurement
    Q0 preliminary result >109
      not much multipacting, cavity and coupler very quiet during last tests
    Lorenz force detuning, cavity tuning sensitivity
      lost some motor steps during 1st movement, others ok
  test of new electronics for cold tuner system as being developed by our Polish colleagues
    preliminary satisfied with the results
Final status/results
  cavity back to CEA, modulator/klystron to ESS
  results presented at SLHiPP-8 (12 June 2018)
  full report published: urn:nbn:se:uu:diva-371627

## Slide 5: Spoke Valve Box Prototype (2019)
Cryo test run with simulator (Dec + Jan)
  commissioning controls & functionality test
Thermo-acoustic oscillations
  installed RLC-circuit to damp the oscillations (test with cryomodule cool down)

## Slide 6: Spoke Cryomodule Prototype (2019)
Warm RF conditioning
  ~3 days/cavity
    1st conditioning ~50 hours
    2nd conditioning ~20 h (cross-contamination)
  fully automatic system conditioning system
    commissioned, now implemented to run 24h/day
  MP bands were consistent with HNOSS test
    strength depends on pulse length, 1st/2nd conditioning…
Cool Down
  4h 15min to 4.5 K
    2h 30 min for shield
Cold RF conditioning (4K)
  no activity
    ~5h both couplers
next step RF test (2K) ...
  cavity conditioning

## Slide 7: Cryomodule Test Plan Draft
VNA
SGD
signal
generator driven
SEL
Lund LLRF
CRYO
Warm Test | Cool Down | Cold Test | Warm-up
Central cavity frequency Qext | Frequency shift vs. temperature | Coupler cold conditioning | Frequency shift vs. temperature
 |  | Cavity conditioning | 
 |  | Central frequency Loaded Q and Qext | 
Coupler warm conditioning |  |  | 
 |  | Steady heat load | 
 |  | Q0 Dynamic heat load Reach 12 MV/m Tuning range of the slow step tuner | 
 |  | Stabilization of the cavity field with LLRF using only RF compensation / with fast piezo | 
NOTE: final optimized & minimized list under discussion (ESS/IPNO/UU)

## Slide 8: Series Spoke Cryomodule Testing (13x)
Acceptance testing of 13 cryomodules
  best estimate: 6 weeks/module
  ESS schedule:
    4 to 6 weeks/module,
    8 weeks for cryomodule S1,
    driven by installation planning
  DESY experience: 1 year to reach “full speed”
Present planning
  first series cryomodule arrives in Summer (4 months delay compared to Dec.2018)
  will “hit” installation schedule by S3 or S4 (transportation!)
ESS SRF schedule dated 2018-12-19
1 year
SRF2015 MOPB080
Yellow bars = planned
Green = real quantity

## Slide 9: Cryomodule Test Schedule
 | time [days] | time [weeks] | in bunker [days]
Arrival, unpacking, initial inspection | 2.5 |  | 
Installation, connection to valve box | 5.0 |  | 5.0
Warm test | 2.0 |  | 2.0
Cool down | 2.0 |  | 2.0
Cold test | 5.0 |  | 5.0
Warm-up* | 4.0 |  | 4.0
Disconnect, packing, shipment | 4.5 |  | 3.0
TOTAL | 25.0 | 5.0 | 21.0
Spare and wasted time, 20% (DESY statistics) | 5.0 |  | 4.2
GRAND TOTAL | 30.0 | 6.0 | 25.2
Time given in 8h work days, 1 shift/day. Not including weekend, holidays and vacation. *) Warm-up is shorter if during weekend |  |  | 
= ~5
weeks

## Slide 10: Foreseen & Unforeseen Issues (1/2)
Operations take longer time than scheduled
  RF conditioning scheduled for 2 days, present indication 3 days/cavity
    investigating conditioning both couplers simultaneous, improving vacuum pumping
No time allotted for re-test or failure
  also: no space at Uppsala to “store” extra cryomodules
Limited staff resources
  No staff for 24h operation
    SSM authorization does not allow un-attended operation if producing radiation
  Not sufficient staff resources for high-peak workload
    difficult to find replacement for high power RF & electronics engineer
    several engineers working part-time (at/near retirement)
  No time allotted for vacation
    university "closure" 5 weeks during Summer, 2 weeks during Winter
    also: multiple sick-leave will squeeze staff resources
  Limited time contract (ESS agreement until Sept. 2020 w/ extension if delay)
    employment rules require dismissal notification one (1) year in advance

## Slide 11: Foreseen & Unforeseen Issues (2/2)
Test equipment failure
  RF station failure
    at present 1 of 2 stations operational
      2 RF stations with 2 tetrodes each + 2 spares = 6 total; at present 3 tetrodes out-of-operation;
      1 tetrode on order, delivery in September; 2 tetrodes back at Thales for investigation
  LLRF/SEL/VNA failure
    2 systems available for each (1 for each cavity)
  Controls failure
    1 system only, some spare parts available but not all
Infrastructure
  electricity, primary cooling water, ventilation from Akademiska Hus
  secondary cooling water, 2 circuits available
  helium liquefier and recovery system
    no spares, but maintenance contracts for compressors
  vacuum
    spares from other systems; 2 ESS pumping carts for beam vacuum

## Slide 12: Summary
FREIA Cryomodule Test Stand up and running !
procedures developed and tested on HNOSS
procedure validation and optimization ongoing on prototype
strongly motivated & flexible personnel to make it work
support from ESS and IPNO when needed (RF/SRF/vacuum/...)
BUT:
tight in scheduled time & available resources