## Slide 1: Lessons Learned from the Ion Source and LEBT Testing and Commissioning - Summary -
Ciprian Plostinar
AD-ICS Testing and Commissioning Manager
European Spallation Source ERIC

## Slide 2: Overview
Topics
  Ion Source and LEBT, Transition from Installation to Testing and Commissioning - Alejandro Garcia Sosa
  Overall Beam Commissioning Experience - Ryoichi Miyamoto
  BI: Diagnostics Deployment - Elena Donegani
  Control System Integration – Saeed Haghtalab
  Software and Machine Studies ICS Perspective – Emanuele Laface
  Software and Machine Studies AD Perspective – Natalia Milas
  Control Room Experience and Operational Perspective – Marc Munoz
  Personnel Safety Systems – Stuart Birch
  Safety and Radiation Protection – Helen Boyer
Summary
  Findings
  Generic Lessons Learned

## Slide 3: Ion Source and LEBT, Transition from Installation to Testing and Commissioning - Alejandro Garcia Sosa
Wrong connections on the coils
Grounding not fit for purpose. Design not suitable for the local conditions at ESS. Overall impact was a delay of ~5 months and damaged equipment.
Cooling water leaks in the magnets.
  Design makes maintenance difficult (hard to access components)
  Water quality issues
Uncomissioned components
  LEBT Chopper.
    Installation setup not suitable for the ESS requirements
    Electrical infrastructure in the tunnel needs improvement
  IRIS
    Different standard for controls
    Wrong cable terminations
Findings

## Slide 4: Ion Source and LEBT, Transition from Installation to Testing and Commissioning - Alejandro Garcia Sosa
Insufficient time for local and integrated tests of each system.
  Rushed process to start commissioning.
  This can be traced down to the fact that for a long time we could not energise equipment.
Undocumented design changes or poorly defined requirements and documentation.
  Swapped magnet polarities. Design changes not documented because it was expected that they would have no impact.
  ISRC Repeller. Parts installed do not match drawings.
  In some cases the polarities weren’t properly defined anywhere.
  Cabling documentation status unclear as each group own their own cables, so coordination is not obvious.
  Issues with water connectors inside vacuum were highlighted during reviews, but not documented.
Lack of common spares.
Lack a comprehensive acceptance test and review of components from IK-partners.
Compartmentalised nature of work at ESS.
  We knew early on that grounding was not done properly, but kept running.
Excessive bureaucracy and volatile rules.
No SAT was held for the Ion Source.
Key personnel lost in the middle of the process. Lack of ownership of various systems.
Generic Lessons

## Slide 5: Overall Beam Commissioning Experience - Ryoichi Miyamoto
Jira plan got behind the reality. Coordination ended up on a short-term basis.
  The plan was too straight-lined and too detailed, and hard to be fixed given all the changes and uncertainties.
We need to define and communicate better mid-term goals.
We should have paid much more attention to technical details of the machine during preparation.
  Many details were missing or not good enough for "end users" (Beam Physics and Operations).
  Communication should be improved and BI and Operations should be involved more during preparations.
  We also spent very little time to get familiarized to interfaces in control room, even for the interlocks.
One useful addition could be a committee that includes line-managers to advise or approve major status changes of systems and subsystems.
  Roles of the line-managers were not necessarily clear, and including them into a committee provides better defined roles.
  This set-up suites the phased approach than one having a single commissioning period
Long-known technical issues were not clearly highlighted in IRR, TRR, nor SRR.
  Solenoid field quality, Chopper design and polarity, Iris vacuum compatibility, Doppler limitation due to its location.
Findings

## Slide 6: Overall Beam Commissioning Experience - Ryoichi Miyamoto
After the commissioning was completed, we believe that most subsystems are “usable”, but not “operational” in a rigorous sense.
  Full testing completed, high-level controls operational, set-points and machine configuration defined, manual available, operator training, etc.
All high-level controls require “cleaning-up”.
We need a much stronger communication on user requirements.
Control Room environment needs improvement.
Findings

## Slide 7: Overall Beam Commissioning Experience - Ryoichi Miyamoto
More effort is needed on supervising and centralising preparation status of each activity.
Transition from Commissioning to Operation should be formulated better.
  We should formulate how to transfer expertise of subsystems (hand-over).
In the absence of MPS, the component protection was done procedurally. This has obvious limitations.
Clarify the use of communication tools: jira, confluence, logbook, meetings, etc.
We need better preparation as too often we had to improvise.
We need a strategy to develop a commissioning and operations team with clear roles and authorities.
Authority and priorities need to be better clarified.
  Beam physics should be involved in the decision making process of start-up.
  Different priorities between line managers and the commissioning team.
Generic Lessons

## Slide 8: Overall Beam Commissioning Experience - Ryoichi Miyamoto
In preparation for and during beam commissioning, we need a forum to highlight technical issues even of we don’t have the resources to tackle them.
We need a separate “beam readiness review”, independent of the SRR since the focus is fundamentally different.
Current status of the Ion Source and LEBT will require more effort to be elevated from ”usable” to ”operational”.
Communication between the field and management can be improved, especially for agreeing on high-level goals and priorities.
Insufficient resources to clean-up identified issues in parallel to maintain the control room activities. This problem is likely to continue in the next phase.
PV naming convention not followed systematically.
For the control room software, the process for approval and handing-over shall be revised, since the process was not working.
Communication process between the developer and user has to be improved.
Generic Lessons

## Slide 9: BI: Diagnostics Deployment - Elena Donegani
Many issues were discovered too late, in the CR
Cable routing and in particular cable termination was a significant issue.
It took a long time to recover from power cuts, both planned and accidental because many systems don’t have automatic rebooting with nominal safe starting parameters so the process is done manually. We had one incident where a planned power cut in one area took out devices in another area that were not supposed to be affected. Automatic rebooting is possible, but it needs to be implemented (BI provides safe working points, ICS does the implementation).
The EMU collided with the Doppler window, this was not identified as a risk prior to installation.
Lack of system replica and LCR replica to anticipate issues in the control room, which means a lot of issues are not discovered before the system is used in the control room.
Lack of timestamp/trigger. PV data have a timestamp that may or may not be possible to correlate to other timestamps (may be off by a small time delta). Having a trigger (integer) number associated with each PV would be extremely helpful for later data analysis (correlation of data).
For the verification without beam, there were a lot of plans (documented in Chess), but many were skipped due to lack of time, lack of manpower, or even lack of power.
Findings

## Slide 10: BI: Diagnostics Deployment - Elena Donegani
The way BI currently split the systems mean they have no redundant system leads. This is again a manpower issue.
We lack a machine overview, we have multiple sources of the machine current status. We need a “status screen.”
A better (more pragmatic) workflow for the OPI development is needed.
Systems replica needed, and assets traceability (Infra has cables, but what is the status?).
In terms of documentation, user manuals, rack cabling documents and device status & history should be documented better.
Findings

## Slide 11: BI: Diagnostics Deployment - Elena Donegani
Preparation and Execution
  Cable routing
  Cable termination
  Assembly verification
  Rack population
  Collision avoidance
  Verification w/o beam
  Operational limits
  Verification with beam
Findings
Interfaces
  OPIs not approved
  OPI workflow needs improvement
  Meeting structure and time can be optimised.
  No alarms
  Interlocks not optimal
  Preliminary ESS Logbook
  Timestamps
  PVs saved manually
  No triggers after EMC

## Slide 12: BI: Diagnostics Deployment - Elena Donegani
Planning and Coordination
  System relocations
  Underestimated EMC
  Lost beam time
  System handover needs improvement and no papers
  Machine status overview
  Public webpages (webpv, etc.)
  Communication with LCR and morning meeting
  Priorities
  No spare time for surprises
Findings
Administration
  No device history (what is the usage integrated time?)
  Archiver
  No spares
  Documentation
  No user manuals
  Database
  Catania legacy
  No asset traceability
Organisational
  No automated tests
  No remote power monitoring
  Limited remote access
  Restart after power cut
  No list of safe/default parameters
  Technicians availability
  No redundant system leads
  No system replica in the lab, no LCR replica.

## Slide 13: BI: Diagnostics Deployment - Elena Donegani
Lack of spares
Documentation, user manuals, etc. needs to be improved.
More testing time is needed to verify systems prior to commissioning
System replica in the lab is essential for debugging, upgrades, etc. to avoid interference with the CR activities.
Asset traceability is critical as we increase the number of devices. Is EAM suitable? Is it widely adopted by all stakeholders?
OPI workflow has to be improved.
High-level overview with the status of each device is needed. One unique source of information would simplify the work. Currently we use many communication channels which are hard to follow: Jira, Confluence, Logbook, Slack, Email, Chess, etc.
The meeting structure and time should be optimised to suit all stakeholders.
Generic Lessons

## Slide 14: Control System Integration – Saeed Haghtalab
Structure, processes and standards were not fully defined since the beginning.
FAT and SAT were not fully conducted and documented.
Documentations (system design, configuration etc.) were not fully completed and reviewed.
We need clarity in role and responsibility distribution (Interlock system, Power supplies, electrical connection and installation, racks, cabling and wiring etc.)
Many valuable experiences and information are lost when key system owners leave.
Many instant decisions and planning in daily operation meetings.
More realistic high level commissioning plan. Move from open loop planning to closed loop with the feedback of stakeholder’s input.
If critical infrastructure activities are significantly delayed, the commissioning project should be re-planned accordingly.
Establish a prioritized list for control systems early in order to allow for correct planning of work.
Inadequate system/integration tests caused many unexpected problems.
Many signs were observed and ignored prior to incidents. (Vacuum gate valve, Coils PS)
Hardware change -> control system redesign -> Cost and time (Chopper, Steerers PS)
Findings

## Slide 15: Control System Integration – Saeed Haghtalab
More established procedure and formality in handing over the systems to operation.
Many of documentations were missing or out-dated.
Operator’s involvement in OPI designs and get in touch directly with software group.
Lack of the controls monitoring tools provided to operators resulted in many calls from LCR to integrators which were not related to integrations.
The high level ISrc-LEBT commissioning planning was not adjusted based on reality resulted in higher cost, effort and time.
Several hardware changes between the Catania control system and ours. There were massive changes on the interlock system.
Findings

## Slide 16: Control System Integration – Saeed Haghtalab
Ensure well-defined requirements and standards and close interaction with in-kind partner prior and after delivery to ensure fully documented configuration and comprehensive well-documented FAT and SAT.
Address commissioning as one cross-division “core team” with clearly defined roles and responsibilities. Engage proactively with the rest of the commissioning team to ensure better coordination of activities.
Define objectives, scopes and deliverables well in advance and revise regularly based on compiled data from everyone involved. Inappropriate planning and schedule pressure would lead to “prototype systems” that are expensive to operate and maintain.
Ensure that all system tests and integration tests are carried out completely prior to starting commissioning activity. This would help system owners to find faults and bugs before which leads to less incidents and system changes.
Ensure that system documentations are well addressed in planning and time estimation.
Issues should be reported to relevant stakeholders and only after initial analysis from operators. This requires good tools and services provided to operators.
The boundary of system ownerships should be clearly defined to avoid overlap in activities.
Generic Lessons

## Slide 17: Software and Machine Studies ICS Perspective – Emanuele Laface
Porting Catania’s IOC to ESS
  IOCs not compatible with E3, so they will have to change again. Naming convention not always followed. Device configuration not always ready for production usage. The infrastructure for creation and deployment only partially used to create the IOCs.
Timing system
  This is a temporary setup.
Network and data infrastructure
  Several issues with EPICS gateway were fixed by removing them. Long term data storage not defined.
CSS
  Significant performance issues. Running in Eclipse is one factor, but also the OPIs that were quickly ported from Catania generate many errors. No OPI responsible.
Online Model simulator and applications
  People prefer to use Python as creation of new applications requires training and time.
Archiver
  It cannot handle big datasets. It cannot be used to store large PVs. It is hard to keep track of what is archived as everyone can add or remove PVs.
Findings

## Slide 18: Software and Machine Studies ICS Perspective – Emanuele Laface
Logbook
  Not final
Python environment
  Stability and usability issues. Many upgrades during commissioning. Some configurations lost at each restart.
  Bandwidth and disk usage and stability issues.
Alarms
  Alarm system is in place, but almost no IOC has the alarms set.
Save and restore.
  Heavy and unusable. Not final.
Findings

## Slide 19: Software and Machine Studies ICS Perspective – Emanuele Laface
Time schedule. We underestimated the effort required so we didn’t have a proper detailed plan and we were not ready for a production setup. The result was a constantly moving schedule and plans with beam time used for hardware local and integrated testing.
Communication. System owners sent very few requirements to the software group in preparation for commissioning about the tools and applications needed. As a result, an overload of requests arrived during commissioning for most devices. We did not have a dry run to identify these needs. We spent very little time in Catania during their tests.
Responsibilities. We require clear owners and responsibilities to avoid the difficulties we had with CSS and orphan OPIs. The Catania control system didn’t have an ICS owner when it arrived.
Priorities. The importance of beam commissioning needs to be reflected in the resources and weight management allocates for this. We cannot lose weeks of beam time for water pipe cleaning that could be done at a later stage. HR was also largely unprepared for night shift work.
Findings

## Slide 20: Software and Machine Studies ICS Perspective – Emanuele Laface
Generic Lessons
We need a strict, but realistic plan, with wide visibility that everyone follows. All stakeholders must understand that their work is critical.
AD should keep ICS in the loop about all devices that need even the simplest control. ICS also has to be proactive in establishing communication channels.
Every system from hardware to software should have a clear person responsible from the end of manufacturing to end of commissioning that ends when handed over in a stable state to operations.
We need clear statements from management at all levels that once beam commissioning has started, it has high priority.

## Slide 21: Software and Machine Studies AD Perspective – Natalia Milas
No time for debugging applications before commissioning. Many fixes had to be done on-the-fly
The Java interface and GUI and the lack of specialised skills, made feature implementation and bug fixing slow. Professional Java coding support would be useful.
No dedicated developer was available full time. People developing the GUI were also commissioning the machine.
OpenXAL updates were cumbersome mainly because it was released as a single big .jar file making it hard to do small updates in the Applications
During commissioning, it was found that the scripting environment was significantly more agile, and completely took over in the commissioning.
There was no clear definition of what polarity of the solenoids should be, or if it even matters (as long as the two follow the same convention). The equipment had no indication of what convention it followed.
This is also true for the correctors.
Diagnostics were not ready and providing wrong readings, making some apps unusable and hard to debug.
Findings

## Slide 22: Software and Machine Studies AD Perspective – Natalia Milas
NPM setup was lengthy, application confusing and hard to use, horizontal reading still has the wrong polarity.
It was challenging to use some of the OPI’s from diagnostics because they needed a lot of configuration every time they were started.
Timing issues.
No clear way to save big data (format and storage) or waveform data. We ended up using a folder mounted in Nextcloud and rely only on the users to put the data back on the correct folder (with dates). Even then it is very hard to retrieve data now since there was no established format and requirement for metadata.
Data. Storage solution not clear. Formats not defined. Wrong data stored. Archiver issues.
EPICS. PV ownership, PV definitions, naming.
PV naming convention not followed systematically and hard to be understood by people other than system owners.
Findings

## Slide 23: Software and Machine Studies AD Perspective – Natalia Milas
Generic Lessons
Primary issues, organisational (lack of manpower), preparations and planning were difficult because there was no time to test connecting to equipment before the commissioning, and equipment became available throughout the commissioning.
Professional Java coding support would be useful.
For the next stage, it is planned to primarily stay with scripts initially, and then develop graphical interfaces once the logic is more mature.
For the next phase Beam Physics should get involved more in understanding available measurements and hardware details for equipment arriving from IK partners.
Conventions have to be understood in advance and enforced. For everything, from Magnets polarities (and have checks prior to start of commissioning) to PVs naming and Archiving setups for the important PV
We need basic measurement devices available at ESS for quick checks.
Diagnostics integration. Dedicated time is needed before commissioning for debugging and testing.

## Slide 24: Control Room Experience and Operational Perspective – Marc Munoz
Issues with lack of spares which are cheap, significantly delaying the commissioning activities for something that only cost a few euros.
The temporary shift leaders did a very good job, but now we have the issue of transition their experience to the incoming operations team that is currently being hired.
Getting data from the archiver is currently somewhat challenging, which makes running analysis more difficult. Looking at the data stored, there is a clear improvement when we look at the data from the end of this commissioning run compared to the start.
One of the most important data for the archiver, the average current during the (3 ms) plateau, is not available. Also, more generally, data saving is cumbersome.
There were no replacement in case of absence of shift leaders, this should not happen in the future (requires more staff)
It is important to track remaining minor issues and fix them before the next commissioning phase.
The issue with the gate valve shows the importance of defining proper procedures (for bypassing interlocks in this example)
Findings

## Slide 25: Control Room Experience and Operational Perspective – Marc Munoz
CS-Studio (with Display Builder) is too unstable and slow to be used when going forward
We need some PV that summarize the machine status and the uptime
Many systems that are considered working will change for the next phase (i.e. timing system)
Elog is acceptable but has some bugs.
We need a clear plan for alarm configuration, logging and operator training.
Save/Restore works but the backend can be improved
Findings

## Slide 26: Control Room Experience and Operational Perspective – Marc Munoz
Generic Lessons
Operations procedures will need to be updated, based on the lessons learned during this commissioning phase. Some procedures will need to be updated during commissioning as well.
The handover procedure needs to be refined, taking in account things not fully working.
We need dry runs and testing without beam before start of commissioning.
We need to address the problems of coordinating and prioritising commissioning and installation.
Priorities need to be clearly set by the management.
Lack of cheap spares is costly.
Change of interlock logic during commissioning has to be done under strict control and with a clear procedure for bypassing.
Clarify and speed the approval of OPIs. Create operator OPIs with only relevant information and well defined layout. What we have for ISrc and LEBT is not scalable for the full NCL.
Manpower clearly needs to be more adequate for the next phase.

## Slide 27: Personnel Safety Systems – Stuart Birch
To produce documents for a PSS system takes a large amount of time and resource. It was the first PSS at ESS and we underestimated the true impact of producing documents.
Document approval too many reviewers and approvers so it took considerable time to get documents finally approved
We spent too much time revising the documents and changing after comments from reviewers.
At the start of the process there was a lack of hazard and risk assessment information. So it was difficult to obtain a real set of PSS0 overall safety requirements.
PSS0 Design software planning and design and simulation took longer than anticipated.
PSS0 Installation was slow. Required improved communication and planning with CF and AD stakeholders. In PSS0 case, late delivery of mains power and Safety fence cost the PSS team a considerable amount of installation downtime.
ESS as a facility was not ready for SRR.
The PSG team felt that the systems that we interfaced with had not been fully commissioned or commissioning had been rushed, HV power supply etc etc. Final integration tests were delayed.
Findings

## Slide 28: Personnel Safety Systems – Stuart Birch
Generic Lessons
Ensure that all stakeholder operational procedure documentation is completed and approved well before operation starts. (key handling, alarm handling, system re-set etc. etc.). Allocate sufficient time for documentation
Improve delivery of commissioned stakeholder interface systems.
It is really important that the hazard and risk assessments are carried out as early as possible
SRR should not be rushed

## Slide 29: Safety and Radiation Protection – Helen Boyer
There was one near miss incident, where beam was on during exchange of radiation monitors inside shielding wall.
Part of G01 found pitch-black
Water leakage during the relocation of the LEBT Faraday cup
Nitrogen gas bottle not properly fixed in a safe state next to equipment
Sparking of cable trays
Masking of interlock signal during integration test of PSS0 for the Ion Source and LEBT Test stand
Temporary removal of shielding material and attached dosimeters
Control Room air quality was lacking
Findings

## Slide 30: Safety and Radiation Protection – Helen Boyer
Generic Lessons
There needs to be a proper management system in place for removable shielding.
Communication practices have to be enhanced
Demonstration of shutdown of equipment if emergency services intervention required
Ensure design and implementation of grounding fully in place

## Slide 31: Overall Positives
Good team of professionals, many dedicated people in the team who go far beyond the expected performance. We have good expertise in house.
Dependencies among subsystems and tuning activities were well understood.
  This helped re-planning while many subsystems were not available.
We were very careful to avoid beam induced damages.
Participating the test in Catania was extremely helpful.
We managed to bring most subsystems up and running
We reacted well to different "situations", e.g., fixing subsystems.
We were reasonably good at improvising, e.g., developing a script on the fly, adjusting the team, etc.
We accumulated a lot of experience.
Good experience with work request process, risk assessment. Collaboration with the rigging team was also a positive experience.

## Slide 32: Overall Positives
Good on-call support
ISrc, LEBT and associated control systems delivered as a consistent and tested unit.
Good collaboration with in-kind partner in a constructive and friendly manner.
The AD-ICS forum was a very big step forward. It reduces the burden on integrators shoulders by defining the priorities and schedules between stakeholders themselves instead of at integrators or WP level.
Vacuum and PSS were well prepared and executed.
Effective team-work with high level of commitment.
Good collaboration and commitment between the people involved in the commissioning.
The dynamics between ICS, Beam Dynamics, diagnostics and Linac groups was quite nice.
Fast fixes and patches (for the correctors control for example)
Support was always available
The ISrc itself works relatively well, it is easy to start and it is stable
Training provided to operators was adequate

## Slide 33: Overall Positives
The PSG team delivered a well structured set of personnel safety system documentation. Excellent system delivery: Design, Installation, Verification and commissioning.
Radiation Protection. Communication with RPO in daily meetings was very good with good support during surveys and dosimeter exchange. Good control of area, good fencing and management of area. ATEX area was respected.

## Slide 34: Overall Lessons Learned
Allow sufficient time for local and integrated hardware testing prior to commissioning with beam.
Improve level of documentation. Ensure requirements and standards are well defined. Allocate time in the planning for the development of documentation.
Make provisions for common spares, consumables and basic measurement devices for quick checks. Where appropriate, develop system replicas in the lab.
Ensure suitable asset traceability.
Hold adequate technical reviews (SAT, TRR). Hold a technical ”beam readiness review” independent of the SRR.
Clarify authority and priorities. Define a commissioning team with clear roles. Assign clear owners for every system from hardware to software.
Improve communication with stakeholders and clarify the use of communication tools. Management should make clear what are the priorities.
Establish and review procedures when testing and commissioning components without protection functions in place.
Allocate sufficient resources to address identified issues during testing and commissioning.

## Slide 35: Overall Lessons Learned
Ensure that processes, conventions, deployments are clear to all stakeholders.
Improve procedure and formality in handing over the systems to operation.
Improve interaction with in-kind partner prior and after delivery of equipment to ensure fully documented configuration and comprehensive FAT and SAT.
Ensure that critical infrastructure is available prior to testing and commissioning. Clarify status on an ongoing basis.
Develop a strict, but realistic plan, with wide visibility that everyone follows. Define objectives, scopes and deliverables well in advance and revise regularly based on compiled data from stakeholders.
Improve the communication and collaboration between AD and ICS.
The commissioning team needs to become familiar with all relevant hardware details prior to commissioning.
Update operation procedures, based on the learning during this commissioning phase.
Change of interlock logic during commissioning has to be done under strict control and with a clear procedure for bypassing.
Improve the SRR process.

## Slide 36: Identified Risks
Loss of key personnel at the wrong time
Testing and commissioning procedures prone to human error
Transition from commissioning to operations not clear/not defined