## Slide 1: IS&LEBT Commissioning Review Safety and Radiation Protection
Helen Boyer / Sigrid Kozielski
ESH Division
European Spallation Source ERIC
28 August 2019

## Slide 2: Positive
Radiation Protection
Communication with RPO in daily meetings
Support during surveys and dosimeter exchange
OHS
Good control of area, good fencing and management of area.
ATEX area was respected

## Slide 3: Negative
Radiation Protection
Management of removable shielding
Near miss (see other slide)
OHS
See fire and electrical issues

## Slide 4: Safety and RP Incidents
ID | Occurrence title | Occurrence unit | Type of occurrence
2019-83 | Part of accelerator tunnel found pitchblack | G01 Accelerator Tunnel (AT) | Risk observation
2019-22 | Beam left on while working inside shielding wall | Normal Conducting Linac (NCL) | Near miss
2019-20 | Water leakage during the relocation of the LEBT Faraday cup | G01 Accelerator Tunnel (AT) | Near miss
2018-170 | Nitrogen gas bottle not properly fixed in a safe state next to equipment | Front End Building (FEB) Level 90 | Risk observation
2018-90 | Sparking of Cable Tray | Normal Conducting Linac (NCL) | Near miss
2018-89 | Masking of interlock signal during integration test of PSS0 for the Ion Source and LEBT Test stand | Front End Building (FEB) Level 90 | Risk observation
2018-80 | Temporary removal of shielding material and attached dosimeters | Normal Conducting Linac (NCL) | Risk observation

## Slide 5: RP Incidents
Beam left on during exchange of radiation monitors
Initial Cause: failure in communication
Root cause: stress applied to the shift leaders to extract as much beam as possible and causing safety measures to be set aside
Recommendations:
Local panel showing if beam is on or off
RP risk assessment of placement of dosimeters
Review communication practices

## Slide 6: OHS Issues
Electrical
Ensure design and implementation of grounding fully in place
Issues associated with high voltage test
Demonstration of safe shutdown of equipment in the event of a need for emergency intervention
Fire and explosion
ATEX assessment and follow up actions – permanent hydrogen monitor
Minimising fire load in area
Shutdown of fire suppression system and no compensatory measures in place
Sealing of pipes and cables between fire compartments

## Slide 7: OHS Issues - 2
Control room air quality
Meeting to understand issues
Checks of air handling units and filters
Deep clean
Change of furniture
Consultant to carry out more detailed analysis

## Slide 8: Top 3 issues
Control of removable shielding
Review of ATEX controls
Demonstration of shutdown of equipment if emergency services intervention required

## Slide 9: Questions