## Slide 1
F. Lafont
J-C. Buffet, S. Cuccaro, B. Guerard, R. Hall-Wilton, K. Kanaki
A new neutron monitor for high flux neutron beams

## Slide 2
ESS Requirements for next-gen monitors
High Flux
High rate
Strong Gamma environment
Minimal material budget and perturbations of the beam
Absolute flux accuracy
Stable over time (< 5 % variations)

## Slide 3
Uranium monitors
Pros
Cons
Best neutron/gamma discrimination
Fast signals / high counting rate
Good stability over time
One major supplier
Bad uniformity over the active area
Efficiency specifications not reliable
High beam perturbations / attenuation
Low voltage operation
Long delivery time
After-sales service
Radioactive sources regulations

## Slide 4
Uranium monitors
Uniformity issues
Measured with a collimated (2.5 mm ×2.5 mm) 2.5 Å- neutron beam
Chart: Axial Uniformity of 926 (Mod. 3053)
| -50 | -47.5 | -45 | -40 | -30 | -20 | -10 | 0 | 10 | 20 | 30 | 40 | 50 | 52.5 | 55
Hor. Diameter | 492 | 4101 | 5735 | 6049 | 6605 | 7030 | 6842 | 7167 | 7326 | 7278 | 6922 | 6415 | 5894 | 3635 | 496
vert. Diameter | 512 | 823 | 5087 | 6481 | 7041 | 7162 | 7458 | 7167 | 7030 | 6828 | 6418 | 5996 | 4921 | 823
Distance (cm)
Distance (cm)

## Slide 5
Uranium monitors
Uniformity issues
Chart: Uniformity of 825 (Mod. 3054)
| -57.5 | -55 | -52.5 | -50 | -47.5 | -45 | -42.5 | -40 | -30 | -20 | -10 | 0 | 10 | 20 | 30 | 40 | 42.5 | 45 | 47.5 | 50 | 52.5 | 55 | 57.5 | 60 | 62.5
 | 20 | 101 | 148 | 761 | 1985 | 2149 | 2306 | 2243 | 2372 | 2740 | 2742 | 2718 | 2792 | 2608 | 2580 | 2332 | 2186 | 2211 | 2103 | 2029 | 683 | 149 | 170 | 46 | 15
Uniformity : up to 20-25 % deviations
Active area : ≈ 100 mm (108 mm in specs.)
Measured transparency : 89 - 96 %
  Average 91% (12 LND monitors tested )
Chart: Uranium 235 total mass (mg)
| 5 | 0.5 | 13 | 1.2999999999999999E-2 | 50 | 1.2999999999999999E-2 | 13 | 13 | 0.13 | 0.5
 | 0.819 | 0.82964 | 0.89908 | 0.93301 | 0.90398 | 0.91601 | 0.92812 | 0.95708 | 0.94203 | 0.87262
 | 0.94208 | 0.93778
 | 0.8982 | 0.92044
  Even worse with a low efficiency one (

## Slide 6
ILL proposal
A multitube gas-filled monitor
Initially designed for SHARP instrument for operation in vacuum
Additional features compared to Uranium monitors
Efficiency can be tuned
Excellent uniformity
Better transparency (0.5 mm windows / 0.4 mm inner walls)

## Slide 7
Multitube Monitor
Pulse height spectra
Chart: Plateau SHARP_V0
| 700 | 725 | 750 | 775 | 800 | 825 | 850 | 875 | 900 | 925 | 950 | 975 | 1000
 | 45.3102 | 50.7456 | 56.1656 | 59.2377 | 60.584 | 60.7744 | 60.7027 | 61.2462 | 61.3838 | 62.308 | 64.2372 | 66.8729 | 70.9736
 | 66.8497 | 64.17 | 62.392 | 61.5602 | 61.0942 | 60.8444 | 60.7862 | 60.5803 | 58.7972 | 56.1634 | 51.6845 | 45.8912
825 V
Chart: Plateau SHARP_V1
| 200 | 250 | 300 | 325 | 350 | 375 | 400 | 425 | 450 | 475 | 500 | 525 | 550 | 575 | 600 | 625 | 650 | 675 | 700 | 725 | 750 | 775 | 800
 | 1635 | 3138 | 4854 | 5094 | 5231 | 5178 | 5182 | 5172 | 5282 | 5375 | 5915 | 7107 | 8787 | 11879
2 bar
4 mbar
  (also tested with Nitrogen)
500 mbar -
10 mbar
Channel

## Slide 8
Multitube Monitor
Gamma discrimination
Beam On
Beam On with shielding
Beam Off
500 mbar
    Channel
    Amplitude (V)
2 bar
  Excellent gamma discrimination @ 2 bar
    Even better @ 500 mbar

## Slide 9
Multitube Monitor
Stability
P.H. Spectra over time
One day of 100s-runs
During Shutdown
Decay
Amplitude (V)
60 kHz on over a day
On H523/SuperSUN @ ILL
  Good stability – No charging effect
  No shift on PH spectra

## Slide 10
Multitube Monitor
Stability
Two day of measurements with 3 monitors
Rate (cps)
Beam intensity variations
Time (s)
Time (s)
Stability curves
Rate (cps)
Time (s)
a.u.
Sharp_V1 counts normalized to Uranium monitors
  ≃ 0.2 % counting rate variations (stat. err.
U – monitor 1
Sharp_V1
U – monitor 2

## Slide 11
Multitube Monitor
Efficiency and Transparency
  Calibration with Hexacomb detector
    Absolute flux
    Attenuation rate by the monitor (scattered + absorbed)
  Efficiency measured to be as expected (then, used for beam flux measurement on the instrument Panther @ ILL)
  Transparency measurement : (expected 97,8 %)

## Slide 12
Multitube Monitor
Uniformity
Chart: Horizontal scan – slit parallel to the tubes
| -3.5 | -2.5 | -2.25 | -2 | -1.75 | -1.5 | -1.25 | -1 | -0.75 | -0.5 | 0 | 0.5 | 1 | 2 | 3 | 4 | 5 | 5.5 | 5.75 | 6 | 6.25 | 6.5 | 6.75 | 7 | 7.25 | 7.5 | 8 | 9 | 10
 | 4954 | 4681 | 4084 | 3169 | 2602 | 2539 | 3053 | 3840 | 4459 | 4759 | 4916.8333 | 4987 | 4886 | 4967 | 4984 | 4990 | 4858 | 4077 | 3079 | 2620 | 2602 | 3314 | 3923 | 4521 | 4613 | 4787 | 4812 | 4912 | 4902
Chart: Horizontal scan – Slit perp. to the tubes
| -5 | -4 | -3 | -2 | -1 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 13037 | 13021 | 12867 | 12981 | 12997 | 13099 | 12950 | 12970 | 12879 | 12949 | 12923 | 12979 | 12929 | 12984 | 13032 | 13056
Chart: Vertical scan – slit perp. to the tubes
| 40 | 41 | 42 | 43 | 44 | 45 | 50 | 55 | 60 | 65 | 70
 | 13189 | 13116 | 13277 | 12992 | 13205 | 13092 | 13144 | 13154 | 13202 | 13240 | 13363
  In the normal case of use with slit apertures, uniformity is great
  (std. dev. < 0.5 % hor. / < 0.8 % vert.)
Beam collimated with a slit (0.5 mm x 12 mm)

## Slide 13
Multitube Monitor
Uniformity
Beam collimated with a mask (20 mm x 20 mm)
  Efficiency mapping within the active area shows a very uniform detector all over the surface
    Statistical uncertainty : 0.4 %
    Std deviation : 0.6 %

## Slide 14
Multitube Monitor
Optional 1D or 2D localization
    This monitor could also be filled so that it allows absolute beam flux measurements (as hexacomb detector but with larger and more channels)
  More measurements still needed to evaluate the accuracy

## Slide 15
Conclusions
Good neutron/gamma discrimination
Fast signals / high counting rate
Optional 1D-2D localization and beam absorption profile
Tested on instrument with strong gamma background
Excellent uniformity
Efficiency from
Small beam perturbations / attenuation
Tested at 60 kHz - 300 Hz/--- Achievable 250 kHz ?

## Slide 16
Thank you for your attention

## Slide 17
Back-up slides
Decay after shutdown
Amplitude (V)

## Slide 18
Backup slides
Pulse height spectra structure
2.5 mm slit centered on tube
2.5 mm slit centered on detector (Al wall)
2.5 mm slit perpendicular to the tubes
Structure observed in the spectrum seems to be due to the presence of Al walls

## Slide 19
Backup slides
Signals