## Slide 1: Science Directorate Instrument Technologies division Neutron Chopper Group TAP 5-2014 Requirements of Instrument suite
Iain Sutton
Neutron Chopper Group Leader
www.europeanspallationsource.se
23 April, 2014
Notes: Good morning
My name is Iain Sutton
I am currently the Group leader for Neutron choppers
And In the following 10mn I will attempt to present the Status and strategy of NG Tech work package

## Slide 2: A preliminary review of chopper systems included in current proposals
Proposal round 2013
Notes: Turning aside from the current preparations for construction I would like to share with you a brief look at some of chopper systems might look like

## Slide 3: Spallation source instruments
PHOTO ISIS Instrument
- overlay schematic components
CHOPPER 3
CHOPPER 1
CHOPPER 2
NEUTRON OPTICS
- GUIDE
NEUTRON OPTICS
- POLARIZER
NEUTRON OPTICS
- COLLIMATION
DETECTOR
-ARRAY
DETECTOR
-BEAM MONITOR
DETECTOR
-BEAM MONITOR
FLIGHT TANK
BEAM LINE SHIELDING
SAMPLE
ENVIRONMENT
DETECTOR
-BEAM MONITOR
NEUTRON OPTICS
- IN MONOLITH GUIDE
BEAM SHUTTER
-LIGHT
BEAM SHUTTER
-HEAVY
Cut away of LET on TS2
Courtesy of ISIS
Notes: ----- Meeting Notes (2/26/14 12:36) -----
The basic anatomy of a neutron spectrometer at a spalation source
Suffice to say that choppers are one of the principal components of most of them and that is unlikely to change at ESS

## Slide 4
BEAM PORT
# 1-2 3-4
BW SELECTION
FUNCTION
NOTES
NMX
GUIDE
2PA-02 (FTF)
14Hz, D700
L 20 50 156
2PA-02 (FTF)
BW SELECTION
14Hz, D700
GUIDE
GUIDE
Status: P1-P.Design
Chopper system schematic
EQUIP ACCESS ZONE
CLASS : BLACK
EQUIP ACCESS ZONE
CLASS : GREEN
Sector: W
Beam-port: W12
Notes: Wide range of designs and system complexities
Ranging from No choppers to more than 20
These are simple ‘stick representations’ and I do not propose to analyse them at length
But flick though a number to give you an over view
Source on the left
Sample on the right

## Slide 5
2PA-01-M-350 (CR)
BEAM PORT
GUIDE
GUIDE
# 1-2 3-4 5 6-7 8-9
PULSE SHAPING
FUNCTION
NOTES
CS/ VOR
Phase 0: Proposal
Schematic
PA-02 (-M-14)
GUIDE
MONOCHROMATOR
PHASE(D) / CO
RRM
BAND WIDTH
2PA-01-M-350 (CR)
2PA-01-M-!$) (FTF)
2PA-01-M-400(CR)
COMMON SHIELDING
SHUTTER
SHUTTER
POSSIBLE SHUTTER POSITION
POSSIBLE SHUTTER POSITION
L 8 16-16.05
Sector: Beam port: NA

## Slide 6
BEAM PORT
7-14Hz, D700
# 1-2 3 -4 5 6-7 8 9-10 11
WAVELENGTH BAND DEFINING
FUNCTION
NOTES
RF/ FREIA v2.0
GUIDE
FRAME OVERLAP
L 6.5/6.6 7.1/7.5 8,5 10/10.1 11.1 15 15,6 25
2PA-02 (FTF)
PA-02 (DC)
2PA-02 (FTF)
PA-02 (DC)
GUIDE
GUIDE
GUIDE
WAVELENGTH BAND DEFINING
WAVELENGTH BAND DEFINING
FRAME OVERLAP
FRAME OVERLAP
PULSE SHAPING
7-14Hz, D1300
7-14Hz, D1300
56Hz, D1300
56Hz, D1300
42Hz, D1300
28Hz, D1500
EQUIPT ACCESS CLASS : ORANGE
EQUIPT ACCESS CLASS : GREEN
GUIDE
2PA-02 (UO)
2PA-02 (FTF)
PA-02 (DC)
GUIDE
GUIDE
Status: P0-Proposal
Schematic
Sector: E/N Beam-port: ?

## Slide 7
BEAM PORT
PA-02 (DC)
GUIDE
GUIDE
280Hz, D700
# 1 2 3 4 5 6 7
PULSE SHAPING (CT1)
FUNCTION
NOTES
PD/ HEIMDAL
Status: P0-Proposal
Schematic
2PA-02 (FTF)
GUIDE
FRAME OVERLAP (CT3)
L 6,5 7 9 13 78 78,6 S = 169m
PULSE SELECTION (CT2)
PA-02 (DC)
14Hz, D1360
14Hz, D700
PA-02 (DC)
BAND DEFINITION (CC1)
14Hz, D1280
FRAME OVERLAP (CC3)
14Hz, D1280
PA-02 (DC)
FRAME OVERLAP (CC3)
14Hz, D1280
EQUIPT ACCESS CLASS : BLACK
PA-02 (DC)
EQUIPT ACCESS CLASS : TBD
Notes: The typical number of choppers on a concept is 7.5
This figure is 3 time higher than the average of instruments at reactor sources and twice that of short pulse spallation sources
The is nothing surprising in this, it is a result of techniques required to get the most out of a long pulse

## Slide 8
BEAM PORT
# 1-2 3-4 5 6-7 8-9 10-11 12-13 14-15 16-17 18-19
PPS
FUNCTION
NOTES
IM-ODIN
28Hz, D700
L 6,5 7 9,25 10 11 12 16 20 30 40 60
2PA-02 (FTF)
BAND WIDTH
GUIDE
GUIDE
PA-10 (PPSc)
2PA-02 (FTF)
2PA-02 (FTF)
2PA-02 (UO)
2PA-02 (UO)
PULSE SHAPING
PULSE SHAPING
56Hz, D700
BAND WIDTH
56Hz, D700
3,5-14Hz, D700
3,5-14Hz, D700
2PA-02 (FTF)
2PA-02 (FTF)
2PA-02 (FTF)
SUB FRAME OVERLAP
SUB FRAME OVERLAP
SUB FRAME OVERLAP
SUB FRAME OVERLAP
14-42Hz, D700-2000
14-42Hz, D700-2000
2PA-02 (FTF)
SUB FRAME OVERLAP
14-42Hz, D700-2000
14-42Hz, D700-2000
14-42Hz, D700-2000
GUIDE
GUIDE
GUIDE
GUIDE
GUIDE
GUIDE
Status: P1-P.Design
Sector: S
Beam-port: S12
Schematic
EQUIP ACCESS ZONE
CLASS : BLACK
EQUIP ACCESS ZONE
CLASS : GREEN

## Slide 9: SUMMARY

## Slide 10
A Summary of Choppers for 22 Instruments, based on 3+16 instrument concepts
22 Instruments = 150 – 160
Notes: Average choppers 7.5
Which 22 inst 150 -160 axis
By configuration, performance and requirements can be sub divided into 6 broad families
Small rotored disc chopper
Low / intermediate / high speed
Large rotored choppers

## Slide 11: Summary
Average 7.5 Choppers per beam line.
2 > 3 times more than Short pulse or Reactor sources
Disc choppers
  40 high speed disc choppers
  Max rotor speed 400+Hz
  70 Low speed disc choppers
  30 Choppers with a rotor diameter of 2m +
  3m diameter rotor under consideration
Prompt pulse suppression
  3 or 4 instruments
Fermi Choppers
  None !
NEW
  Fan Chopper

## Slide 12: Chopper performance
PA-2-H-C
PE-10-V-M
PA-5-H-S
PA-1-H-M
PE-5-V-M
Notes: We looked around the world at chopper systems …

## Slide 13: A suite of Choppers for ESS
Preliminary estimations

## Slide 15
A Suite of Choppers… for 16 Instruments,
based on current instrument concepts
16 Instruments ~ 120 axis
LOW
SPEED
(48)
PA-02 (-M-14)
(20)
INTERMEDIATE
SPEED
PA-02 (-M-14)
HIGH
SPEED
(18)
2PA-02 (UO)
LARGE
ROTOR
(24)
2PA-02 (FTF)
(3)
FAN
PPSc
(6)
PA-10 (PPSc)
Notes: Average choppers 7.5
Which 22 inst 150 -160 axis
By configuration, performance and requirements can be sub divided into 6 broad families
Small rotored disc chopper
Low / intermediate / high speed
Large rotored choppers

## Slide 16: ‘Tendency’
No ‘Fermi’ Choppers
  Except Tempis fugit ‘rotating mono array’
The ‘Fan Chopper’
  Several groups now expressing interest…
Requests for 3m rotors is declining
Request for a ‘compact’ PPS chopper
Requests for ‘movable’ choppers
  Variable separation
  Out of beam

## Slide 17: Platform 1 Disc chopper – Small rotor – Low speed
DC-SR
(50-70)
Key requirements
Rotation speed: Low ( 7 - 96 Hz)
Openings: Large
Attenuation at short wavelengths
High reliability in radiation environment
Low lifetime cost
Principal characteristics
Type: PA-1-H- (horizontal axis disc)
Rotors : Simple + Robust
Diameter: 600-1000mm
Enabling Technologies .
Rotors materials: Metallic
Bearings: Magnetic

## Slide 18: Platform 2 Disc chopper – Small rotor – Intermediate speed
DC-SR
(15-35)
Key requirements
Rotation speed: (96-192 Hz)
Openings: (Multiple) Small or Large
High reliability in radiation environment
Principal characteristics
Type: PA-1-H- (horizontal axis disc chopper)
Rotors: Simple or Optimized
Diameter: 600-800mm
Enabling Technologies .
Rotor material: CFRP / Alu
Bearings: Magnetic

## Slide 19: Platform 3: Disc chopper – Small rotor – High speed
Key requirements
Rotation speed: 192- 400+ Hz
Openings: (Multiple) Small
Minimal guide interruption
High reliability in radiation environment
Principal characteristics
Type : PA-1-H- (horizontal axis disc)
Rotors design: Optimised
Diameter : 600-700mm
Enabling Technologies .
Rotor material: CFRP / Ti / MMC
Bearings: Magnetic
DC-SR-VHP
(25-35)
2.PA-1-H-M

## Slide 20: Platform 4: Disc chopper – Large rotor
DC-LR
(25-35)
Key requirements
Rotation speed: (7 - 56 Hz)
Openings: Multiple V.Large, Asymmetric
High closing speed
High reliability in radiation environment
Principal characteristics
Type: PA-1-H- (horizontal axis disc chopper)
Rotors: Optimized
Diameter: 1200 - 2000mm
Enabling Technologies .
Rotor material: CFRP / Alu
Bearings: Magnetic or Contact

## Slide 21: Platform 5: Disc chopper – Fan
DC-SR-FAN
(3-5)
Key requirements
Rotation speed: (7 - 56 Hz)
Openings: Variable width, Asymmetric
Principal characteristics
Type: XPA-1-H- (horizontal axis disc chopper)
Rotors: multiple concentric ‘Fan’
Diameter: 500 - 800mm
Enabling Technologies .
Rotor material: Aluminium
Bearings: Contact

## Slide 22: Platform 6: Prompt pulse suppression chopper PPSc
PPSc
(4-8)
PA-10 (PPSc)
Key requirements
Rotation speed: (7 - 56 Hz)
Closure: 3ms, Symmetric or Asymmetric
Attenuation : 90% @ Prompt pulse energies
Extreme radiation resistance
Principal characteristics
Type: PA-10-H- (horizontal axis chopper)
Rotors: 300 – 400 thick , single or double
Diameter: 500 - 600mm
Enabling Technologies .
Rotor material: Nickel alloy / Tungsten
Bearings: Magnetic

## Slide 23: Platform 56 ‘Instrument specific developments’
Support the development of ‘one off, instrument specific choppers’ which can not be achieved within platform approach.
Support in integration into ESS systems
  Control
  Support systems
  Pit systems
Active collaboration in engineering design process and decisions.
Examples
  PST
  Compact PPSc

## Slide 24: MEETING FACILITY GOALS
Best service – Least cost
Reliability/Availability/Serviceability
Life time cost

## Slide 25: Balancing act
  Aggressive operating schedule
  (225 beam days)
  Complex instruments
  (130+ choppers)
  Availability > 95% (?)
Notes: ESS Instrument availability target = 95%
(Availability = Beam time – Down time)
  This is very ambitious for a spallation source and comparable to the best reactors.
Target mean time to repair 24hrs IMPORTANT POINT!
  Very challenging, will require 24/7 technical support and full parts backup on site.

## Slide 26: Instrument suite model
INSTRUMENT 01
<-------------- 15 degree ------------- >
INSTRUMENT 03
< 5 degree >
INSTRUMENT 02
TARGET MONOLITH
Notes:   Beam lines were divided up into 2m long ‘modules’ from R = 2 to R = 24.
  Conditions required for access considered with Proton beam on target.
  Consider mechanical components only
  Optical and steady state components are assumed to be reliable ie MTBF >100 x Mechanical components

## Slide 27: Serviceability
Lateral shielding
Access levels
Mean Time To Repair (MTTR)
Time to access (zone dependent)
Time to repair (equipment dependent)
Notes: Evaluate for each 2m section of each beam-line
Quantity & Type of equipment installed
Level of accessibility (p-beam ON)
Access time

## Slide 28: SNS Beam line/Group totals & availability – FY13 thru June 2, 2013
“SNS neutron beamline/instrument availability is about 89%”
Chart: 
 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 1 | 1
 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 1
 | 3 | 1 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 3
Beam interruptions
Cycles (Normalized to 35 day)
CAUSE OF STOPPAGE
BLUE Chopper system
RED Other component
GREEN At request of I.S.
~ 3 ‘significant’ interruption per 100 beam days
+3.5 uses of shutters to effect ‘running repairs’
Instrument suite downtime Data from ISIS, TS1 & 2 (2007-2011)
Instrument suite downtime
Notes: As I have said we were interested in modeling reliability
So went out search for data
From the other facilities reactor or spallation sources
ISIS
SNS
J-PARC
ILL
Frequency of shutdowns
Duration
Cause

## Slide 29: RELIABILITY Instrument related causes of beam loss
Source: ISIS
Note: excluding - Sample environments, Computing & Motion control
Causes often requiring access within heavy shielding for repair
Causes not requiring access within heavy shielding for repair
Instrument MTBF ~598 days
Instrument suite MTBF ~35 days
Notes:   Calculated failure rates for existing instruments and equipment
  And from that
  Factoring in the increase comlexity of ESS instruments tried to come up with some estimates for our own instruments
  Failure of a specific instruments every 3 or 4 years
  Failure of ‘An’ instrument each 35day cycle

## Slide 30: Simulated ESS Instrument availability (averaged across suite )
Chart: 
| option A /B | Option C/D | Option E
225 BD | 66.4 | 93.4 | 81
215 BD + 4 Short breaks | 79.8 | 93.4 | 89
Notes: This is what we came up with – the red bars
By tweaking the schedule to improve the distribution of maintainaince periods were able to reduce the time that instruments risk to remain idle following a serious brakdown.
And for the cost of a few days of beamtime we currently have this E
It’s not yet there – there is a shortfall
but what is important is that we are aware of it and can now feed this into the design of instruments …

## Slide 31: ESS Equipment reliability targets
 |  | Working hours |  |  | shutter option |  |  |  |  |  |  | 
Reliability Classification |  | MTBF |  | MTBR | Name |  | Distribution |  | Typical components |  |  | 
Class A |  | 72000 (10-15yrs) |  | 36 | Very reliable - quick fix |  | 70% |  | Std choppers, Vacuum systems, Fixed Collimation, Std optics |  |  | 
Class B |  | 72000 |  | 168 | Very reliable – slow fix |  | 10% |  | Special Choppers Special optics |  |  | 
Class C |  | 3600 (8-12mnths) |  | 36 | Reliable – quick fix |  | 15% |  | diaphragms |  |  | 
Class D |  | 3600 |  | 136 | Reliable – slow fix |  | 5% |  |  |  |  | 
Note |  |  |  |  |  |  |  |  |  |  |  | 
These figures concern only |  |  |  |  |  |  |  |  |  |  |  | 
Equipment installed within the common shielding area (radius 24m) |  |  |  |  |  |  |  |  |  |  |  | 
Equipment presenting a significant risk of breakdown - ie no supports, shielding |  |  |  |  |  |  |  |  |  |  |  | 
Equipment critical to instrument performance - options and non essentials have been excluded |  |  |  |  |  |  |  |  |  |  |  | 
Notes: Estimates of MTBR for shutter option A – F will always be shutdown within 3days in order to effect repairs. This requires that the 'cost of the shutdown' be multiplied by the number of beams shutdown

## Slide 32: MEETING PROJECT GOALS
Requirements from a different set of stakeholders …

## Slide 33: Project Drivers
Raising Performance
  Neutronic performance
  Availability
Reducing Costs
  Procurement
  Operations
In-kind
  Harness experience
  Maximise contribution
  Equitable distribution

## Slide 34: EARLY SUCCESS
Success will be measured early (within 3 years)
Tranche Instruments must be receiving users and publishing major papers by 2021 ….
Priority Trinity
Construction
Commissioning
Early Operations

## Slide 35
 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028
Operations Schedule |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#5 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#6 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#7 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#9 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#10 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#11 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#12 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#13 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#14 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#15 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#16 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Initial Design | Final Design | Procurement & Fabrication | Construction | Installation | Cold Commissioning | Hot Commissioning
1st Neutrons
Instrument projects overview

## Slide 36: Chopper & Instruments The ‘first 3’
Initial Design | Final Design | Procurement & Fabrication | Construction | Installation | Cold Commissioning | Hot Commissioning
1st Neutrons
On site Installation
Delivery, Integration & Test
Procurement
Development
Commissioning Interventions
Target Operations Schedule |  |  |  |  |  | 
#1
#2
#3
T-60m T-48m T-36m T-24m T-12m T0

## Slide 37: HOW MANY & WHEN
Period | Phase | PPS | Disc | Fermi
2013-2016 | Technologies development | 4 | 3 | 1
2017-2019 | Build program ‘Day 1 suite’ | 2 | 35+ | 3
2020-2025 | Build program ‘Full suite’ | 4 | 85+ | 6
2025 > | Total choppers in service | 6 | 125 | 10
2013-2016
2017-2019
2020-2025
Notes: The number are quiet staggering
The total requirements are in excess of that of any existing facility by a factor of around 2 to 3
This is not totally unexpected and many of the techniques which the long pulse makes possible RRM and WFM require a large number of choppers to be employed …
What is sure is that is a whole lot of choppers….. and that we are going to need a lot of help developing solutions and then supplying the need.
We need partners to pull this off both academic and industrial and in order to help us work together effectively we need good organization and a central plan.

## Slide 38: THE END
Thank you for you attention