## Slide 1
Loss monitoring and simulations
A. Jansson
(work by L. Tchelidze)
2012-03-21

## Slide 2: Contents
BLM time response;
Loss distributions and monitor location
BLM total dose;
Detector testing plan
Loss limits for the ESS linac

## Slide 3: BLM time response
Technical Note ESS/AD/0031
  http://eval.esss.lu.se/cgi-bin/DocDB/ShowDocument?docid=168

## Slide 4: Time response conclusions
The amount of time in which the full ESS beam would start melting steel/copper can be 1.5 μs - 25 μs for the beam radius of 1 mm - 4 mm. This time increases sharply as the energy increases.
The BLMs system should be able to generate an abort signal in 2us.

## Slide 5: Loss Distribution
Technical Note ESS/AD/0032
  http://eval.esss.lu.se/cgi-bin/DocDB/ShowDocument?docid=169
Loss location = middle of first quadrupole
Loss angle = 1.5 mrad
Loss intensity = 10^12 protons/sec

## Slide 6: Observations
The maximum dose rates are observed closest to the beam pipe (as expected).
The dose rates before the magnet in which the loss occurred is usually around 10 times smaller than that after the magnet.
If the losses happen in a given inter-cryomodule sector, the secondary particle showers are observed even in the next inter-cryomodule sector (the higher the proton energy, the more comparable the dose rates in the two inter-cryomodule sectors are).
The prompt dose rates outside the cryomodule is 3-4 orders of magnitude lower than those observed close to the loss point.

## Slide 7: Conclusions BLM placements
One ionization chamber before and after each quadrupole magnet. Thus we should use 3 loss detectors per quadrupole doublet.
An additional ionization chamber in the middle of each cryomodule to monitor average beam losses.
Detectors should be placed closest possible to the beam pipe or cryomodule.
Detectors should be placed perpendicular to the beam pipe (applies to the inter-cryomodule sectors). Outside the cryomodules, detectors can be placed longitudinally to the tank.

## Slide 8: BLM total dose
ESS - 5300 hours/year
40 years
Total of ~200 khours
20 GGy total dose to the BLM (for 1 W/m losses)

## Slide 9: Detector testing @MAXLAB
Response time:
  Max-II injection point:
  e- beam, 380 MeV, 800 ns, 10 Hz.
(Gamma) response functions:
  Cs-137 (661 keV), Co-60 (average 1.25 MeV) Na-22 (1.27 MeV), Mn-54 (834 keV).
Simulated SNS response

## Slide 10: Activation limits
Technical Note ESS/AD/0026
  http://eval.esss.lu.se/cgi-bin/DocDB/ShowDocument?docid=163
L. Tchelidze &
J. Stovall