## Slide 1: BPM electronics (An overview of the requirements and the possible solutions for ESS)
Hooman Hassanzadegan
ESS, Beam Instrumentation Group

## Slide 2
BPM system requirements
Platform for the BPM electronics
Overview of the signal detections methods for the BPM pickups
BPM electronics
  Analog front end
  Digital unit
Timing requirements
Next steps
Outline

## Slide 3
BPM system requirements
The total number of BPMs shall be 107 including:
MEBT: 4
DTL: 8 (2 per tank)
Spoke cavities: 30
Low-β elliptical cavities: 10
High-β elliptical cavities: 28
HEBT: 22
The BPM system shall be able to measure:
Beam position
Beam phase
Beam intensity

## Slide 4
BPM system requirements (cont’d)
The BPM pickups will have different designs. Nevertheless, it is preferred to use the same electronics for all the pickups and adapt it to the pickups through minor modifications if needed.
Pickup pictures from SNS, Jim O’Hara, LNAL, “SNS Linac Beam Position Monitor pickups”, 2002

## Slide 5
BPM system requirements (cont’d)
Beam
14 Hz
2.86 ms
Fast measurement
(instantaneous)
Slow measurement
(average)
50 mA
50 mA
+
BPM
Nominal pulse
14 Hz
50 mA
2.86 ms
Diagnostics pulse
1 Hz
50 mA
100 μs
averaging
The BPM system shall be able to make a fast and a slow measurement of the beam position, phase and intensity in nominal and diagnostics pulse modes.

## Slide 6
6
Beam Current
Beam
RF voltage
envelope
14 Hz
Trigger
352 MHz
2.86 ms
RF voltage
Beam position accuracy: ±100 μm (RMS)
Beam position resolution: 20 μm
Phase measurement accuracy: ±1° (RMS)
Phase measurement resolution: 0.2°
Measurement range: 50% of the beam pipe dia.
Phase measurement range: ±180°
ADC sample rate: > 1 MSPS
Electronics response time: < 1 μs
Refresh rate (end user): 14 Hz
We are discussing with the Beam Physics group to finalize the requirements.
BPM system requirements (cont’d)

## Slide 7
Electronics platform comparison
*** Full table by Miha Rescic from Cosylab (Currently working at ESS)
 | VXS | μTCA | PXI | cPCI/cPCIe
Vendor number | 16+ | ~10 | 70+ | Many
Maturity | Medium | Medium | High | High
Max. transfer rate | Star/dual star: 30Gbps, Mesh: 112 Gbps | 1 GbE, 2GbE, 10 GbE; Up to 12.5 Gbps 250MB/s/lane (PCIe) | PXI: 132 MB/s PXIe: 250 MB/s/lane (up to 16 lanes) | PCI: 132 MB/s PCIe: 250 MB/s/lane (up to 16 lanes)
Backplane timing and synchronization | 112 S.E. pins, 32 diff. pairs | 84 pair for MCH, 7 pairs per AMC for 12 AMCs | 10 MHz and 100 MHz clock, 8x shared trigger, star triggers | 
Topology | Star, dual star, full mesh | Star, dual star, full mesh | Master-slave, star | Master-slave, star
Redundancy | YES | YES | Power Supply only | Power Supply only
Users | Military, Avionics | DESY XFEL | EBG MedAustron,… | CERN, LANL, ORNL, ITER (planned), JET, ETE, EBG MedAustron
Crate cost | High (From 3000 €) | Medium (From 2000 €) | Medium (From 880 € for 4 slot version) | Medium (From 900 € for 6 slot version)
Documentation |  | Poor | Fair | Good
Ease of use |  | Fair | Easy to start | Fair
RTM | NO | YES | NO (alternative: FlexRIO) | NO
Power | 90 W (5V), 66W (3.3V) | 20 – 80 W per slot | 30 W per slot | 30 W per slot

## Slide 8
μTCA chassis platforms (ELMA)

## Slide 9
Slide from Bastian Lorbeer’s presentation at the DITANET workshop at CERN, Jan 2012
μTCA systems at DESY

## Slide 10
Simple diode detector and S&H + LPF (LHC – CERN and DESY)
From Marek Gasior’s presentation at the DITANET workshop at CERN, Jan 2012
BPM signal detection (method 1)

## Slide 11
11
BPM signal detection (method 2)
ESS-Bilbao BPM system
Commercial log-amp detector
ESS-Bilbao log-amp detector board
AD8310 response
Signal detection with log-amps (ESS-B and CERN)

## Slide 12
BPM signal detection (method 3)
Signal detection with an IQ demodulator:
9.028 MHz
9.028 MHz
From Claire Simon’s presentation (CEA) at the DITANET workshop at CERN, Jan 2012
IQ demodulator (ESS-Bilbao BPM)

## Slide 13
BPM signal detection (method 4)
Down-conversion to IF + IQ detection (SNS)
Photo ref: J. Power et al. , Beam Position Monitor System for the SNS LINAC, Proc. PAC 2003
+I
+Q
-I
-Q
+I
+Q
-I
-Q
I
Q
f = fIF , Tsampling = n TRF + TRF /4
f = fRF +/- fIF
f = fRF
Mixer
Input
Ref.
IQ (under-) sampling
Demux + sign correction
+I
+Q
From the pickups
Down conversion to IF
ADC sampling
ADC clock generation (PLL)
IQ DEMUX (FPGA)
Calibration timing + AFE gain control (FPGA)

## Slide 14
14
BPM signal detection (method 5)
200
400
40ns
Sample No.
ADC Digits
40ns
Sample No.
ADC Digits
200
400
FPGA calculates envelopes of RFFE signals. Beam position calculated from integrals.
RFFE output, digitized by BPM electronics (5Gsps), after
gain/offset/jitter correction & FIR filtering in FPGA.
5ns
Pickup Voltage
Time
Low-passed pickup signal (scope measurement).
Slide from Boris Keil’s presentation at the DITANET workshop at CERN, Jan 2012
Ringing filter + ADC sampling at very high frequency (SLS):

## Slide 15
Comparison of signal detection methods
 | Diode | Log. Amp. | Homodyne (sampling in Baseband) | Heterodyne (sampling in IF) | Ringing filter + Fast ADC
Complexity | Low | Low | Low | medium | high
Cost | Low | Low | Medium | High | Very high
Performance | Acceptable | Good | Good (with error compensation) | Good | Good
Possibility of use for LLRF | No | No | Yes | Yes | No
Compatible with LLRF
Not Compatible with LLRF

## Slide 16
AFE implementation on μRTM
picture from Bastian Lorbeer’s presentation (DESY BPM) at the DITANET workshop at CERN, Jan 2012
The analog front-end can be developed in house through external collaborations.

## Slide 17
Digital electronics
The digital electronics such as the FPGA unit and the ADC module (COTS) can be directly bought from the market.
ATCA carrier board with RTM and COTS ADC (DESY)
Off-the-shelf AMCs

## Slide 18
Auto-calibration and linearization
R. B. Shurter et. Al, Analog Front-end Electronics for Beam Position Measurement on the Beam Halo Measurement, Proc. PAC’01
J. Power et al, “Beam Position Monitor Systems for the SNS Linac”, PAC’03
Maurice Cohen-Solal, Design, test and calibration of an electrostatic beam position monitor, Ph. Rev. ST - AB
DC offset compensation, linearization and auto-calibration can be done by the FPGA.

## Slide 19
Beam Current
Beam
RF voltage
Envelope
Timing requirements for the BPM system
14 Hz
Trigger
352 MHz
Intra-pulse:
ADC sampling
Data storage on the AMC memory, etc.
Inter-pulse:
Data communication to the control system
Auto calibration
Error compensation, etc.
RF reference phase:
RF phase detection
IQ demodulation
2.86 ms
RF voltage

## Slide 20
LLRF 1
CPU
L.L. elect.
Timing
CPU
Event Gen.
BPM 1
CPU
L.L. elect.
LLRF n
CPU
L.L. elect.
Timing
Fanout
Event Rec.
Ethernet
Coaxial cable
Optical fiber
Optical fiber
Optical fiber
352 MHz
MO
Timing system for BPM and LLRF
Ref:
CDR -> 5. Control Systems
CDR -> 4.8.2 Low Level RF control
HMI
Interface to the control system

## Slide 21
Summary of the minimum timing requirements for the LLRF and BPM systems
Timing signal | Freq. | Synchronization | Spec. | Source | Comment
RF ref. ph. | 352 MHz or 704 MHz | Synchronization needed for the whole Linac | Drift: <0.5° | Coaxial cable or Optical fiber | 
Pulse rate | 14 Hz | Synchronization needed for the whole Linac | Stability: 1ppm | Event based | + Programmable delay
ADC/DAC sampling | 10 MHz – 100 MHz | Synchronization needed within the shelf |  | backplane | 
Non real-time or slow | --- | Not needed | --- | Ethernet | 
Calibration ref. (BPM) | 352 MHz | Not needed |  | Internal to the shelf | 

## Slide 22
Space allocation for the electronics
RF distribution from the klystron gallery to the tunnel, source: ESS RF group

## Slide 23
Next steps
Choosing an appropriate signal detection method
Analog front end prototyping
Tests with a μTCA crate
Building a BPM test bench