## Slide 1
Applications of electron linear accelerators for radiotherapy
Lars Hjorth Præstegaard
Aarhus University Hospital

## Slide 2: Outline
Basic theory of electron linear accelerators (linacs)
Rationale of radiotherapy
Electron linacs for radiotherapy
Exercise: How to change beam energy of linacs for radiotherapy?

## Slide 3: Basic theory of electron linear accelerators

## Slide 4: Linear acceleration of an electron
Requirements for acceleration:
Electric field
Electric force in direction of v
Significant linear acceleration Extended linear overlap of electric field and electron trajectory
Lorentz force:
e: Elementary charge
E: Electric field vector
v: Velocity vector
B: magnetic field vector
F=-eE
v
Electron
E
v
E
E
E
E
Linear electron
trajectory

## Slide 5: Electromagnetic wave in free space
Electromagnetic wave in free space:
Transverse electromagnetic wave (TEM wave)


No electric field opposite to the direction of propagation: No acc.
Both electric and magnetic field perpendicular to direction of wave propagation (z axis)
: Angular frequency (2f).
z: Wavelength
kz : Wave number (2/).
c: Speed of light
Dispersion relation for TEM wave (propagation along z axis) :


Phase velocity ≡ /kz = c
Group velocity ≡ d/dkz = c
≡ 2f = 2c/z
≡ kzc
Dispersion relation: Relation between  and kz (z)
Planar wave:
E(z,t)=E0Exp(i(t+kzz))
wave in +z dir.
wave in -z dir.

## Slide 6: Electromagnetic waves in a waveguide
Waves in a waveguide:
TE: Transverse electric (not suited for particle acceleration)
TM: Transverse magnetic (suited for particle acceleration)
Maxwell equations + Boundary conditions

## Slide 7: Linear acceleration in a waveguide
Generator
frequency
=c (free space)
wave in +z dir.
wave in -z dir.
Stopband
Dispersion relation for TM mode:
 Phase velocity for TM mode = /kz > c

No average transfer of energy to electrons


## Slide 8: Disc-loaded waveguide
 Reflections from discs
Disc-loaded waveguide: Addition of discs to the waveguide:
 Positive interference of reflections from neighbor disks if
Disk-loaded waveguide with period d
Phase advance = kz*2d  2  kz  /d
 Group velocity close to zero for kz  /d
 Standing wave behavior for kz  /d (no energy transport)
 Large perturbation of waveguide dispersion relation for kz  /d

## Slide 9: Disk-loaded waveguide: Dispersion
Dispersion relation for disk-loaded waveguide:
wave in +z dir.
wave in -z dir.
Passband
Disks  Frequency exist for which = c
 Acceleration of relativistic electrons (v=c)
/d
kz

## Slide 10: Disk-loaded waveguide: Modes
Loss-free propagation: kzd = 2/p

p=1,2,3,...
Modes used for particle acceleration:
0 mode: kzd=2 (p=1)
 mode: kzd= (p=2)
2/3 mode: kzd=2/3 (p=3)
/2 mode: kzd=/2 (p=4)
Small holes in discs
Positive interference: phase advance = kz*pd = 2

p=2
Scattering of forward wave (hole=source of wave propagation)


## Slide 11: Traveling wave (TW) acceleration
Disk-loaded waveguide (slowed-down wave)
TM wave and electron beam travels synchronous
Input of RF power at first cell
Output of RF power at last cell
Injection of electron beam along axis of waveguide
RF load
electron gun
Resistive loss in walls
+ Energy transfer to beam

Reduction of microwave power along waveguide
RF power in

## Slide 12: TW acceleration: Stanford linear accelerator
50 GeV electrons
932 disc-loaded sections of 3.05 m
Water cooling
RF input
Discs

## Slide 13: Standing wave (SW) acceleration
2. Low wave attenuation along structure
1. Disk-loaded waveguide with reduced apertures at ends:
Standing waves
Acceleration by both forward/backwards waves (0, -mode)
d
 Full reflection of traveling waves at structure ends

## Slide 14: SW acceleration: /2-mode
-mode: Large energy gain
Short fill time (large group velocity)
Insensitive to geometrical errors
Low energy gain
/2-mode:
/2 mode
bi-periodic /2 mode
Coupling cavity
Biperiodic /2-mode SW structure:
All advantages for /2 and  modes
Looks like -mode for beam

## Slide 15: SW acceleration: Medical linac
Varian 600c biperiodic /2-mode SW structure:
RF input
coupling cavity
Normal cavity

## Slide 16: SW acceleration: Medical linac
Varian TrueBeam biperiodic /2-mode SW structure:
Coupling cavity

## Slide 17: Rationale for radiotherapy
18

## Slide 18: What is radiotherapy?
~50 % of all cancer patients receives radiotherapy
Indirect DNA damage
Direct DNA damage
Radiotherapy:
Damage to DNA by ionizing radiation:
Radiotherapy:
Killing of cancer cells by ionizing radiation (x-rays or ionizing particles)
Dead of cell at cell division

## Slide 19: Why does radiotherapy work?
Cancer cells are sensitive to ionizing radiation:
Cancer cells are mutated cells with reduced DNA repair capability
Cancer divide (copying of DNA) more than healthy tissue + Higher sensitivity to radiation at cell division
Before treatment
After treatment
During treatment

## Slide 20: Dose response
Tumor control probability
Normal tissue complication probability
100 %
0 %
Dose
Response
Dose: Compromise between cure and toxicity to healthy tissue

## Slide 21: Electron linear accelerators for radiotherapy

## Slide 22: Main components of medical linac
Microwave
amplifier
Electron gun
RF waveguide
Disc-loaded
waveguide for acceleration
Bending magnet
Treatment head
Gantry
Couch

## Slide 23: Main manufacturers of medical linacs
Varian
Elekta
4-22 MeV TW accelerator
4-22 MeV SW accelerator

## Slide 24: Gantry and couch degrees of freedom
Gantry
Couch
Isocenter:
Intersection of gantry rotational axis and collimator rotational axis
(100 cm from x-ray target)
Collimator

## Slide 25: Treatment head
Disc-loaded
waveguide
X-ray treatment: Fast electrons + target  Intense bremsstrahlung (x-rays)
Electron treatment: Fast electrons (no target)
Target
Flattening
filters
Bending
magnet
Secondary
collimators
Multi-leaf collimator
Dual monitor chamber

## Slide 26: Bending magnet
Elekta
Varian / Siemens HE
Varian (low energy)
Large beam spot
Achromatic
deflection
Achromatic
deflection
Chromatic
deflection

## Slide 27: Bending magnet: Varian Clinac HE
Energy slit:
Slit at position with non-zero dispersion
Target

## Slide 28: X-ray target
X-ray target:
Located inside vacuum
Conversion of electron beam to bremsstrahlung (x-ray)  X-ray treatment
Target materials:
Target materials affects x-ray yield and spectrum
Copper/water for cooling
Varian Clinac x-ray target:
6 MeV bremsstrahlung spectrum:

## Slide 29: Primary collimator
Primary collimator:
Large tungsten block defining the maximum treatment field size
Effective shielding of leakage radiation
Usually opening with a conical shape  Round maximum field size

## Slide 30: X-ray flattening filter
Bremsstrahlung is forward-peaked
Convenient with flat dose profile
Low energy
High energy
Varian flattening filters:
 Use flattening filter:
Cancellation of off-axis intensity variation
Reduction in dose rate
Off-axis photon energy spectrum variation

## Slide 31: Monitor chamber
TrueBeam dual transmission monitor chamber
Determination of treatment beam dose
Two chambers: Redundant dose determination
Dual transmission ionization chamber:

## Slide 32: Light field
Light field = extend of treatment field

## Slide 33: Secondary collimators (or jaws)
Varian secondary collimators (tungsten):
Secondary
collimators

## Slide 34: Multi-leaf collimator (MLC)
MLC:
2 rows of thin tungsten blades
Detailed shaping of the treatment field
Typical leaf width: 5 mm
Varian 120 leaf MLC (leaf width: 5 mm)
MLC

## Slide 35: X-rays: Secondary electron cascade



Compton scattering
Bremsstrahlung
Pair
production
Patient
e-
e-
Photoelectric
effect
e-
e-
e+


Annihilation
= Ionization
Low ionization (dose) at patient skin
= Dose buildup
e-
e+
Photoelectric
effect
Photoelectric
effect

## Slide 36: X-ray dose versus depth
Dose buildup
Decreasing dose:
Attenuation
Inverse square law
Skin dose ≈ 25 %  Skin sparing

## Slide 37: Electron treatment
Detailed field shaping:
Lead end-frame cut-out
Elekta electron applicator

## Slide 38: Electron dose versus depth
Ionization tracks (20 MeV electrons):
Dose
Depth
Bremsstrahlung tail
Water
Electrons are used for cancer close to the skin
Electron range depends on electron energy
High ionization (dose) at surface

## Slide 39: Example: 5 field prostate x-ray treatment
Field 1
Field 2
Field 3
Field 4
Field 5
Overlap of all fields at cancer target
 Large dose in cancer target relative to dose in healthy tissue
Prostate target
Transverse view of patient pelvis

## Slide 40: Example: Intensity-modulated RT (IMRT)
IMRT: Modulation of each field with MLC  Dose distribution fits better to the cancer target

## Slide 41: Imaging systems on-board accelerator
kV source on
robotic arm
MV radiation
head
Imaging systems:
2D MV imaging
2D KV imaging (good contrast)
2D KV imaging + gantry rotation: CT scanning of pt. in treatment position
MV detector on
robotic arm
KV detector on
robotic arm
Imaging systems  Verification of patient position

## Slide 42: Video
Elekta medical accelerator

## Slide 43: Exercise: How to change beam energy of linac for radiotherapy?

## Slide 44: Change of RF input power
Change of RF input power
Change of the electric field in the disc-loaded waveguide
Change of the energy, capture efficiency, and energy spread
Small energy spread
Capture efficiency: 44%

Large percentage of electron reach the target
Large energy spread
Capture efficiency: 36%

Small percentage of electrons reach the target

Low dose rate + stray rad.
Only optimum capture efficiency and energy spread for a particular RF input power (beam energy)
 Problem for multi-energy linac  Design compromise
Siemens linac: 6 MeV
Siemens linac: 18 MeV

## Slide 45: Change of RF frequency
t1
t2
RF frequency shift  Change of phase velocity 
Desynchronization of wave and particle  Particle energy decrease
Elekta Synergy
Only small change of electric field in the buncher section
Design of optimum capture efficiency and energy spread for a wide range of beam energies.
High dose rate for all treatment energies
Buncher

## Slide 46: Change of number of active cells
Motorized energy switch:
Modification of cavity coupling
Reduced electric field
Same electric field in buncher for all beam energies
 Design of optimum capture efficiency and energy spread for wide range of beam energies
 Efficient transfer of electrons from gun to target for all energies
 High dose rate for all energies
Varian Clinac
Buncher
+RF input

## Slide 47: Appendix: Treatment workflow

## Slide 48: Step 1: Patient fixation
Creation of patient-specific fixation.
Patient fixation:
  High geometrical accuracy.
  High positional reproducibility.
Head & neck fixation
Fixation for breast treatment

## Slide 49: Step 2: CT scanning
CT scanner: Acquisition of 3D patient anatomy for treatment planning.
CT coordinate system is marked on patient or fixation using room lasers:

## Slide 50: Step 3: Target delineation
Delineation of:
Gross tumor (visible tumor).
Suspected microscopic tumor spread (CTV).
Critical healthy tissue
Dose:
Specification of the dose to all targets.
Target delineation is performed by a radiation oncologist

## Slide 51: Step 3: Target delineation
Target delineation is the largest uncertainty in modern high-precision radiotherapy:
Large improvement of target delineation with PET imaging
Lung tumor
Delineation by several radiation oncologists

## Slide 52: Step 4: Treatment planning
Aim:
Get specified dose to all targets (<5 % accuracy required).
Get as low dose as possible to critical healthy tissue (organ specific).
How?
Choice of modality (x-ray or electrons).
Choice of beam directions (gantry angle and couch angle).
  The more beams the better sparing of critical tissue.
Beam weight (dose), shaping (MLC) and modulation (MLC).

## Slide 53: Step 5: Treatment
Position the patient in patient-specific fixation
Adjustment of the patient position using room lasers and marks on the patient/fixation
Correct the position of the patient position using on-board kV or MV image systems