## Slide 1
Medical application of
proton accelerators
Lars Hjorth Præstegaard
Aarhus University Hospital

## Slide 2: Outline
Basic theory of the proton cyclotrons
Production of medical radionuclides
Rationale for proton therapy
Treatment delivery of proton therapy

## Slide 3: Basic theory of proton cyclotrons

## Slide 4: Cyclotron: The basic concept
Magnetic field

## Slide 5: Circular orbit
Circular orbit:
m: Mass of ion.
q: Number of charges
e: Elementary charge
v: Speed of ion
r: Radial position of ion
Bz: Magnetic field strength in the z direction
Angular frequency:
Gyrofrequency or cyclotron frequency
Typically 10-30 MHz
Centrifugal force = Magnetic Lorentz force

z

## Slide 6: Isochronous cyclotron
 Bz(r) = (r)B(0) ,
Problem:
Reduced orbit separation  More difficult extraction
Isochronous cyclotron:

 Field increases versus r + r is limited
CW high intensity proton beam
Simultaneous acceleration of low/high energy protons
Revolution period should be energy independent (isochronous cyclotron)

## Slide 7: Isochronous cyclotron
Methods for increasing the field versus r:
Decreasing pole gab versus r
Trimming coils
Increase hill/valley ratio
Isochronism: The magnet poles must be machined and shimmed with very high accuracy (relative field error below 10-4)
IBA C235

## Slide 8: Transverse focusing
Betatron tune: Number of oscillation periods per revolution
Cylindrical symmetric magnet pole:
Radial betatron tune: r2=1-n
Vertical betatron tune: z2=n
Weak focusing: Less than one oscillation per revolution
Strong focusing: More than one oscillation per revolution
No beam loss  Both radial and vertical focusing
r2=1-n > 0 and z2=n > 0
0<n<1 (weak focusing)
Small field decrease versus r

## Slide 9: Vertical focusing
Vertical defocusing for isochronous cyclotron
Vertical Lorentz force:
=0 for field with no dependence on 
0 for field with dependence on 
Additional vertical focusing: Non-cylindrical magnet pole
Problem:
Isochronous cyclotron: Magnetic field versus r
Vertical focusing: Magnetic field versus r
Solution:
Increasing field versus r + additional vertical focusing
Not
compatible!

## Slide 10: Vertical focusing: Azimuthally varying Field
Azimuthally-varying field (AVF):
AVF  B0:
AVF + spiral pole geometry:
Additional vertical focusing
Isochronous cyclotron with strong transverse focusing

## Slide 11: Vertical focusing: IBA medical cyclotron

## Slide 12: Synchro-cyclotron
Synchro-cyclotron (non-isochronous cyclotron):
B versus r (vertical focusing)
Relativistic ions: m(r) increases versus r

Only acceleration of a limited number of proton bunches (neighbor orbits) per acceleration cycle
Low duty factor (significantly reduced beam current)
decreases versus r
Very compact cyclotron  Very high magnetic field (>5 T)
 Significant saturation of iron
Significantly reduced azimuthal field variation
Insufficient vertical focusing for isochronous cyclotron

## Slide 13: Synchro-cyclotron: Acceleration cycle
IBA S2C2:

## Slide 14: Synchro-cyclotron: Example
Mevion synchro-cyclotron for proton therapy (250 MeV, ~9 T):

## Slide 15: Production of medical radionuclides

## Slide 16: Cyclotron for radionuclide production
RF input
RF input
RF cavity
RF cavity
GE PETtrace 700
Spiral sectors
Targets for radionuclide production

## Slide 17: Cyclotron-produced radionuclides
Acceleration of protons or deuterium by a cyclotron

## Slide 18: Positron emission tomography (PET)
511 keV
511 keV

## Slide 19: Rationale for proton therapy
17

## Slide 20: Interactions of a proton in matter
1. Inelastic Coulomb interaction with atomic electrons
Dominating interaction:
Ionization (=dose)
Small energy loss per interaction  Continuous slowing down of proton  Well-defined range
Range secondary electrons < 1mm  Dose is absorbed locally
No significant deflection of protons (mp = 1832me)
2. Elastic coulomb scattering with nucleus
3. Non-elastic nuclear interaction
Lateral scattering of treatment field versus depth
Neutrons are the main external radiation hazard

## Slide 21: Energy loss of a charged particle
Energy loss per unit length (stopping power):
To first order: -dE/dx  1/v2
NA: Avogadro’s number, re: Classical electron radius, me: Electron mass, z: Charge of the projectile, Z: Atomic number of the absorbing material, A: Atomic weight of the absorbing material, c: Speed of light, v: Velocity of the projectile, β=v/c, γ = (1 − β2)−1/2, I: Mean excitation potential of the absorbing material, δ: Density correction arising from the electron shielding, C: Shell correction item
200 MeV protons in water
200 MeV protons in water

## Slide 22: The Bragg peak
Dose = Number of protons  stopping power  1/v2
 Most dose is deposited at the final range of the charged particle
Dose beyond peak: Dose from inelastic nuclear interactions
Width of peak: Range straggling
Position of peak: Given by the initial particle energy

## Slide 23: Spread-out Bragg peak
Less dose
Less dose

## Slide 24: Dose distribution of proton treatment
Proton therapy: Less dose to healthy tissue
1. Less complications (same dose)
2. Better tumor control (higher dose,
~10 % of patients can significantly benefit from proton therapy
same complications)
Conventional x-ray
Advanced x-ray
Advanced proton

## Slide 25: Treatment delivery of proton therapy

## Slide 26: Synchrotron for proton therapy
Special features:
Proton linac
Variable-frequency RF cavity
Slow extraction

## Slide 27: Synchrotron: Slow extraction
Spill cyclus:
Fill ring with ~109 protons
Accelerate (~0.5-1 s)
Slow extraction during 1-10 s
Decelerate (~0.5-1 s)
Slow extraction:
Betatron tune close to resonance  Reduced area of transverse separatrix
Transverse RF excitation of the beam
 Particles excited outside separatrix
 Particles spiral outside extraction septum
Proton therapy  Need for CW beam

## Slide 28: Accelerators types for proton therapy
 | Cyclotron | Synchrotron | Synchro-cyclotron
Magnetic field | Fixed/1-2 T | Vary/1-2 T | Fixed/5-9 T
Particle radius | Vary | Fixed | Vary
RF frequency | Fixed | Vary | Vary
Focusing | Strong (AVF cyclotron) | Strong | Weak
Size | Compact | Large | Very compact
Energy | Fixed | Vary | Fixed
Beam pattern | Continuous beam | Pulsed | Pulsed
Beam current | High | Low | Low

## Slide 29: Cyclotron: Energy change
Energy slit:
Selection of treatment energy
Energy
slit
Dipole
Dipole
Degrader
Apertures
Degrader:
Beam energy 
Beam size and divergence 
Apertures
Cyclotron
Degrader + ESS transmission:
Courtesy by PSI

## Slide 30: Cyclotron-based proton therapy facility
Degrader: Energy adjustment
Energy selection system (ESS)
Gantry
(100-200 tons)
Beam line
Cyclotron

## Slide 31: Dose deposition in the patient
Passive scattering:
Pencil beam scanning (PBS):
No requirement of patient-specific collimators and compensators
Interplay effect for moving organs
Simple
Straight forward treatment of moving organs
Excess dose to normal tissue
Significant neutron dose
Patient-specific collimators and compensators
Dose distribution of raw beam:

## Slide 32: Gantry
Varian gantry in Munich
Proton gantry:
Weight: 100-200 t
Diameter: ~10 m
Speed: 360/min
Treatment center accuracy: < 1 mm

## Slide 33: Treatment room
Varian ProBeam:
Gantry nozzle
Patient couch
Imaging
system

## Slide 34: Proton therapy vender videos
IBA proton therapy
Varian proton therapy
Mevion proton therapy

## Slide 35: Challenges in proton therapy

## Slide 36: Lateral proton scattering in the patient
Lateral scattering of protons in water:
Beam size in air:
~4-6 mm (energy dependent)

## Slide 37: High sensitivity range uncertainties
Causes of range uncertainty:
CT number to stopping power conversion (200 mm depth: uncertainty of ~6 mm)
Tumor shrinkage
Organ motion during treatment (respiration, rectum gas, bladder filling etc.)
Consequence: Large treatment margin :=(