## Slide 1
Multi-Blade tests @ BNC/WIGNER
Preliminary results
Francesco Piscitelli
ESS Detector Group
22 - March - 2016

## Slide 2
neutrons
5 degrees
θ
Efficiency 45% at 2.5Å
A single Boron layer inclined at 5 degrees
Calculated Efficiency
The intensity is spread over a wider surface
(5 degrees = factor x10)
substrate
10B4C
The Multi-Blade detector

## Slide 3
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
The Multi-Blade detector

## Slide 4
neutrons
strips Y
wires X
substrate
10B4C
cassette
The Multi-Blade detector

## Slide 5
neutrons
strips Y
wires X
substrate
10B4C
Electric field modeling

## Slide 6
neutrons
neutrons
strips Y
wires X
Old Design New Design
cassette
cassette
substrate
10B4C
strips Y
wires X
10B4C
substrate
Materials evaluation

## Slide 7
Any material holding the strips at 5deg scatters too much!
neutrons
cassette
strips Y
wires X
10B4C
substrate
Old Design
scattered or absorbed flux
Materials evaluation

## Slide 8
Efficiency saturated (5degrees)
At 5deg not much survives to the 10B4C
Substrate as thick as necessary
Kapton (strips) shielded by 10B4C
neutrons
strips Y
wires X
cassette
substrate
10B4C
New Design
Materials evaluation

## Slide 9
Multi-Blade mechanical design
amplifier board
amplifier board
cathode strips
anode wires
holding frame
10B4C substrate
neutrons
10x10cm2 demonstrator

## Slide 10
amplifier board
amplifier board
cathode strips
anode wires
holding frame
10B4C substrate
10B4C substrate
holding frame
cathode strips
cathode strips

## Slide 11
10B4C substrate
holding frame
cathode strips
cathode strips

## Slide 13
cathode strips
substrate
anode wires
sputtered 10B4C layer
x-axis (wires)
y-axis (strips)

## Slide 14
140mm
100mm
The Multi-Blade detector
9 Cassettes inclined at 5 degrees with respect to the neutron beam
1 Individual readout (32 wires+32 strips) + 8 charge division (32 channels)
neutrons
neutrons
Multi-Blade

## Slide 15
Uniformity and overlap
neutrons
strips Y
wires X
substrate
10B4C
uniformity
overlap
shadow
shadow

## Slide 16
Uniformity and overlap
neutrons
strips Y
wires X
substrate
10B4C
uniformity
overlap
shadow
Edge (knife) cassette N
cassette N+1
shadow
cassette N

## Slide 17
Knife gain corrected
Knife gain not corrected
neutrons
strips Y
wires X
substrate
10B4C
uniformity
overlap
shadow
Uniformity and overlap
Individual readout

## Slide 18
Spatial resolution X – wires

## Slide 19
Spatial resolution X – wires

## Slide 20
Spatial resolution X – wires

## Slide 21
Spatial resolution X – wires
Step 0.25mm
Wire 3
Wire 2
Wire 1
Single Wire Algorithm
Mutual Information
Resolution FWHM
~ 0.61mm
Resolution FWHM
~ 0.67mm
CoG Algorithm

## Slide 22
Spatial resolution Y – strips
Step 2mm
Step 4mm
Non-linear at edges
(expected and due to parasitic C of strips)
Cassette center
MAX Algorithm
SUM Algorithm
Cassette center
Cassette edge
Cassette edge
Step 4mm
Step 2mm

## Slide 23
FWHM ~ 2.4mm
Spatial resolution Y – strips
Step 4mm
Step 2mm

## Slide 24
Uniformity Y – strips
Cassette edge
Cassette center

## Slide 25
Stability

## Slide 26
Efficiency
~55% at 4.1Å, ~65% @ 5.1Å (5 degrees)
measured
calculated
PHS

## Slide 27
PHS and Gain
Gas Gain ~ 10 - 20
Energy Resolution ~50 KeV
Energy Threshold ~100 KeV
@ 4.1Å 5 degrees
measured
calculated

## Slide 28
PHS
Event 1
Event 2
Wire 11 12 13 14 15 16 17
When the alpha escapes the layer almost orthogonally, it hits the opposite cassette and does not release the full energy.
Integral is the same,
just higher energies shifted to lower energies.

## Slide 29
Rise time
Time to go from 10% to 80%
Time to go from 10% to 90%
No shaper – Pre-amplifier signals
PHS
PHS

## Slide 30
Spot
Counting rate capability
No space charge saturation observed up to ~ 2kHz/mm2
Slit
If we assume the same charge of the slit concentrated in one mm2, rate is above 40KHz/mm2. (Attention: It is not exactly the same!)
Rate per mm2
Rate in the slit

## Slide 31
Images

## Slide 32
Images

## Slide 33
Images

## Slide 34
Images

## Slide 35
Summary
 | Multi-Blade 2013 | Multi-Blade 2015 (actual status) |  | Goal | What to do next
Uniformity X | 15% | 20% (measured in Charge Div. 1350V) | - | <5% | Mech. Improvements and new meas. at lower gain 800V and individual r/o
Uniformity Y | 15% | 20% (measured in Charge Div. 1350V) | - | <5% | Mech. Improvements and new meas. at lower gain 800V and individual r/o
Overlap | 2mm gap 50% loss | 0.7mm gap 50% loss | o | ≤1xResoltution (≤0.5mm) | Mech. Improvements and new meas.
Spatial Res. X | 0.3mm | 0.6mm (wire pitch can be adjusted) | + | 0.5mm | Nothing – Matches the goal
Spatial Res. Y | 4mm | 2.5mm | + | 2-3mm | Nothing – Matches the goal
Stability | Not measured | ~1% (over 12h) | + | 1% (over days) | Longer tests
Efficiency | Measured 26% @ 10deg 2.5Å Extrapolated 43% @ 5deg 2.5Å | >44% @ 2.5Å Measured 55% @ 5deg 4.1Å Measured 65% @ 5deg 5.1Å | + | >40% @ 2.5Å | Nothing – Matches the goal
Low gain operation | Gain ~60 | Gain ~20 | + | -- | Nothing – Matches the goal
Counting Rate Cap. | Not measured (dead time limited by electr. 2us) | > 2 KHz/mm2 (>40KHz/mm2 extrapolated) | o | 500 KHz/mm2 (4 MHz/mm2) | Needs to be measured
Gamma sensitivity | Not measured | Not measured (but full eff. with 100KeV Threshold) | o | <10-6 | Needs to be measured, but in principle is already good due to the threshold used, not different from what measured on other prototypes.
Dynamic range | Not measured | Not measured | o | 109 - 100 | Needs to be measured
Self Scattering | Not measured | Not measured | o | ≤10-4 | Needs to be measured
Fast neutrons sensitivity | Not measured | Not measured | o | <10-6 | Needs to be measured
Reproduce scientific output | Not measured | Not measured | o | -- | Needs to be measured