## Slide 1: Utilities and Infrastructure D01/D03 IKON 11
Clara Lopez – Loki & FREIA Integration Engineer
www.europeanspallationsource.se
Lund, September 2016

## Slide 2: Topics
EPL Buildings/areas
Gallery location and section
Infrastructure details [1]
Cave
Path ways
LoKI and Freia overview in D02/D03
Civil construction [2]
[1] Instrument projects – CF Information- https://ess-ics.atlassian.net/wiki/display/SPD/CF+Info
[2] ESS-0066249 – Status Civil construction – Division meeting

## Slide 3: EPL – ESS Plan Layout
D03
D01
Main entrance
Accelerator
E01
E02
D02
G02

## Slide 4: Instrument Halls D01/D03
D03
Bay door
Nitrogen filling station
Forklift access
Personnel access
Lab D04
D01
D02

## Slide 5: Buildings in D03
Building Joint
Target building
Instrument hall Building
Substation
51 m
33m

## Slide 6: Buildings in D01
Building Joint
Target building
Instrument hall Building
Labs
Labs
Substation
130 m
56 m

## Slide 7: Building joint/ Heights

## Slide 8: Instrument Halls and labs
8
N
E03
E04
E01
E02
D03
D01
D02
D04
D08
Hall3
W1 NMX
W2 Beer
W3 C-Spec
W4 Bifrost
W5 Miracles
W6 Magic
W7 T-Rex
W8 Heimdal
E03
1 × Engineering
E04
1 × LS/SCM
1 × cold room
1 × instrument room
1 × Basic chemistry
1 × characterisation
Hall 1
E1 Estia
E7 Vespa
E8 Skadi
S2 Odin
S4 Dream
S11 Vor
D07
1 × LS/SCM
1 × cold room
D08
1 × Basic chemistry
Hall 2
N1 HR-NSE
N5 Freia
N7 Loki
D04
1 × LS/SCM
1 × cold room
2 × instrument room
Total = 14 + RML
Courtesy: P.Henry
D07

## Slide 9
Floor capacity and floor level in D01, D03, E01, E02.
Floor settlement requirement under the instruments
The floor in the experimental halls and bunker area will due to dynamics loading have a stability for elastic movement.: 3mm
Stability for creep/deformation over the lifespan of the facility: 3mm

## Slide 10: Non-magnetic details in D03
ESS procedure for material choice in vicinity of NSE instruments, ESS-0036612.

## Slide 11: Routing of infrastructure D03 Access gallery – Level 090
https://ess-ics.atlassian.net/wiki/display/SPD/CF+Info
From the gallery
From conduits

## Slide 12: Routing of infrastructure D01 Access gallery – Level 090
https://ess-ics.atlassian.net/wiki/display/SPD/CF+Info
From the gallery
From conduits

## Slide 13: Gallery section
Preliminary

## Slide 14: Gallery distribution
14
SPARE
NSS-CT
ICS/DMSC-CT
SPARE
SPR-D03
N2G
Maintenance access way
2100x1200
FHV
3000
500
AREA FOR LOCAL GROUPS OF NSS CONNECTION
-
HR
HRR
IAR
DIW
ORW
CWL
CF-CT
DHL
800
AREA FOR PENETRATION
900
2600
Notes: Cooling Water Low temperature CWL
Instrument Air (Compressed Air) Deionized Water Nitrogen Gas FHV Fume Hood Ventilation Helium Recovery Helium Recovery Return Cable Tray Integrated Control System Data Management System Center Ordinary Risk Waste Water Sprinkler D03 Experimental Hall 2

## Slide 15: Conduits detail location

## Slide 16: Nitrogen/Deionized water/Compressed Air/LAN
Nitrogen gas (10bar)
Deionized water (8 Bar)
Compressed air (6Bar)
LAN: 2 outlets/connections
Liquid nitrogen
1 filling point / sector
East and South Sector
14 connections from the galery and 7 on the wall.
North Sector
9 connections from the gallery and 2 on the wall.
Process flow chart: ESS-0046901, Layout: ESS-0046984
East and south sector
14 connections from the gallery and 7 on the wall

## Slide 17: Cooling water
Properties
7Bar, 8°C
Cooling capacity / average instrument: 27kW
Bunker
Separate cooling loops for each sector provided by ESS. Connections are by the beamlines.
(1 skid/sector, dedicated pipes for each instrument)
North Sector
9 connections from the D03 gallery.
East and South Sector (D01)
17 connections from the gallery.

## Slide 18: Fume hoods and exhaust pipes for sample environments
North Sector
5 connections from the D03 gallery.
East and South Sector (D01)
8 connections from the gallery.
ESS-0056058, ESS-0048000, Layout: ESS-0046984

## Slide 19: Tap and waste water
1 drain to radiological waste water system in D03
North Sector
1 water connection in D03 on the wall
1 drain to risk water system in D03
East and South Sector (D01)
2 water connections in D01 on the wall
4 drains to risk water system in E01
Process flow chart: ESS-0046901, Layout: ESS-0055781, ESS-0055780, ESS-0046984

## Slide 20: CF Power outlets
Utilities power for installation and maintenance:
3pcs. 16A 1-phase outlets
1pcs. 16A 3-phase outlet
1pcs. 32A 3-phase outlet
North Sector
1 connection from the galery level 090 and 7 on the wall level 100.
East and South Sector (D01)
2 connections from the gallery level 090 and 20 on the wall level 100.
LAYOUT: ESS-0046984

## Slide 21: Instrument Power distribution Concept
Substation
Gallery/Halls
Instrument distribution panel
/Main grounding zone./
Instrument distribution panel
/Grounding zone X/
Bunker distribution panel
/Bunker zone/
Lights
Motion
...
Chopper
Vacuum
Light
Detector
...
Motion
SAD
Instrument 1
Instrument 2
Instrument 3
Instrument 4
Instrument X
CF responsibity CF budget
NSS responsibity, Instrument budget
NSS responsibity, NSS budget
Preliminary

## Slide 22: Instrument Power distribution Concept
15
38
36
Bunker:
The bunker is a unified zone. There will be four distribution panels by D02 walls for the four sectors.
Cave /Main zone/
There will be connection points similarly placed as the Sample environment connections
ESS Power and Grounding: ESS-0051373 (Not released)
The design will be provided by ESS, the installation and material cost will be paid by the instruments. The design will be finalized according to the instrument designs.
There will be no central UPS system
The power distribution network will be planned according to ESS grounding rules.
All the missing insulation rules and solutions will be defined ASAP
Preliminary

## Slide 23: Cabinet placement
Preliminary
Constrains on cabinet placement:
Cable length limitations
Floorspace
Distance from sensitive components
Blocking of trafic routes including cranes
The cabinets should be outside of the shielding
Bunker:
NSS is planning to have cabinet arrays for each sector in order to minimize the cable lengths. The cabinets will be connected to the power distribution panel of the bunker sector. Also they will be connected to a common ICS router.

## Slide 24: ICS/DMSC
ICS/DMSC network will be provided by ESS. The network design will be finalized according to the instrument designs.

## Slide 25: Cranes
D01 : ESS-0055303 (5ton), ESS-0055308 (7.5ton), ESS-0055305 (30ton)
D03 : ESS-0055307 (5ton), ESS-0055304 (7.5ton), ESS-0055306 (30ton)
E01 : ESS-0052338 (10ton)
E02 : ESS-0052339 (5ton)
E03, E04: ESS-0048477 (5/15 ton)
Mobile crane (preliminary): https://ess-ics.atlassian.net/wiki/display/SPD/Mobile+Cranes+in+E02

## Slide 26: Cranes

## Slide 27: Jib cranes for sample area
LoKi Crane
CF has approved the mounting of the LoKi Crane to D03 floor.
Arm length: 7m
Capacity: 1500kg.
Mounting recommendations and will be release soon.

## Slide 28: Considerations for logistics inside the halls
General logistics: planning of logistics has started, comments are welcome:
https://ess-ics.atlassian.net/wiki/display/SPD/Logistics+in+D01%2C+D03%2C+E01%2C+E02
Crane logistics: Not analyzed in detail
Shielding logistics: Not analyzed in detail
Logistics for vehicle and personal traffic: Not analyzed in detail (,but)

## Slide 29: Additional information
Sprinkles: Technical Specification D01-D08_E01-E05 ESS-0013209
Air conditioning: ESS-0056058, ESS-0048000
Considerations for Free – paths in instrument halls:
NSS Logistics https://ess-ics.atlassian.net/wiki/display/BO/Free+path+by+the+building+wall

## Slide 30: Cave - Loki
Cooling water
Tap water
Power outlets
Gasses supply
Communication
Compressed air
ESS-0038163 Sample environment utility supplies
Supply infrastructure:
Power
Network
Lighting
Ventilation
Fire protection BFS 2011:26
Local crane

## Slide 31: Cave - Freia

## Slide 32: LoKI - Cooling diagram

## Slide 33: Access to sample area
Local crane
Lifting cage (equipment)
Sample area preparation
Pathways
Personnel access
1
2
3
4
5
Required information about recommendations for:
Mounting and Location for local cranes
Grounding practical details
Instrument height limitations
Anchoring system

## Slide 34: Path ways
2550 mm [3]
[3] NSS Logistics https://ess-ics.atlassian.net/wiki/display/BO/Free+path+by+the+building+wall

## Slide 35: FREIA- LoKI overview

## Slide 36: Freia – Loki Bunker wall - Integration

## Slide 37: Integration in bunker wall
According to ESS-0062127
Short instrument wall
Feedthru outline

## Slide 38: ESS Buildings
Personnel access

## Slide 39: Civil construction
Tunnel G01 Outside

## Slide 40: Civil construction
A2T

## Slide 41: Civil construction
TARGET

## Slide 42: ESS Orthophoto August 2016

## Slide 43: Civil construction
TARGET

## Slide 44: Civil construction
Instrument Halls

## Slide 45: Civil construction
Instrument Halls

## Slide 46: Civil construction- reinforcement 20 ton/m2

## Slide 47: Questions?