## Slide 1: Instrument maintenance within the Bunker
28-9-2016
Iain Sutton
On behalf of the Neutron Technologies Division

## Slide 2: Intro
Why is having a coherent maintenance strategy important ?
Because it contributes to;
  Instrument availability for science
  Operating costs
  Minimizing exposure of personnel
  The facilities overall reputation

## Slide 3: Context
ESS specific context
  Long source instruments
    Components@ R6
  Shielding common volume
    Vertical access
HPower source issues
  Neutron activation
  Neutron damage
Project constraints
  Minimise initial investment
Notes: Large open shielding volume reminisant of reactor beamlines ILL FRM2
As a spallation source shielding thicknesses are consequent and it tightly surrounded by instrument
Contrany to rIt does not contain only neutron guides
But Large quantities of mechanical components

## Slide 4: Equipment within Shielding Bunker
Component
Type Numbers
Choppers 98 axis
Optics 300 m (150 sections)
Shutters 30
Collimators 10
Monitors 30
Supports 200
Service supplies
  Vacuum
  Fluids
  Cables
22 instruments assumed (13L-9S)
60-70% of all ‘mechanical’ components on instruments are within bunker volume
2 -3 times more mechanical devices on a ‘long pulse instrument’ than at existing types
Notes: As 2 or 3

## Slide 5: ESS Schedule Steady-State Operation
Cool down period for components inside bunker
No maintenance during beam production !
Shut downs
2 long 30-50days
  Maintenance
18 short shut downs <3days
Notes: Basic constraint
No maintenance on any during beam production
Maintenance will be carried out during shutdowns
Short shutdowns currently considered too short to use – cool off , cranage

## Slide 6: Interventions
Scheduled activities
Type 1A: Periodic
(Once in the life of an instrument ( >10yrs))
  Installation / Decommissioning
Type 1B: Regulars
(Periodicity 2-5yrs)
  Inspection
  re-alignment
  Overhaul (preventive maintenance)
Unscheduled activities
Type 2: Breakdown
INTERVENTIONS
within Bunker
SCHEDULED
Type 1(A)
PERIODIC
Type 1(B)
REGULAR
UNSCHEDULED
Type 2
BREAKDOWN
5-10 interventions p.a.
6 interventions p.a.
60 interventions p.a.
(interventions in steady state operation for 22 instruments
post burn-in)
Notes: Interventions may be considered as two types;
Type 1: Scheduled activities
  Type 1A: Long periodicity events
    Installation / Decommissioning
  Type 1B: Short periodicity events
    Inspection/ Adjustment /Replacement
Type 2: Unscheduled activities
(‘random’ events with 5-20 incidents p.a.)
  Breakdown repair
    Inspection / Adjustment / Replacement
  Accident scenario

## Slide 7: Activation after short cool down period (24 hrs)
Activation
Integrated beam
Beam power
Beam days
Removal of hotspots required prior to access
Personnel access restricted
Personnel access prohibited
MIXED REGIME
Personnel access unrestricted
Personnel access severely restricted

## Slide 8: Activation after cool down period (7-14 days)
Activation
Integrated beam
Beam power
Beam days
Removal of hotspots required prior to access
Personnel access restricted
Personnel access prohibited
MIXED REGIME
Personnel access unrestricted
Cool down
Personnel access severely restricted

## Slide 9: Activation
Time after beam off
Component 24hrs 7days
T0 choppers 10 0.3
Monolith face 1 1
H. shutters 2 0.1
BP shutter 0.1-1 0.05
Collimator (Cu)0.2 0.02
Guide (Al) 1 0.01
Estimated dose at 30cm in mSv
Activation of the monolith face is likely to severely restrict work in proximity
REMOVE or SHIELD

## Slide 10: Dose limits & Zones
0.1 -1 mSv/h
Restricted access
> 1 mSv/h
Strictly controlled access
Zone I
Requirement;
All systems shall be remote handling installation compatible
Zone II
Requirement;
All systems shall be vertical handling installation compatible.
‘Hot’ components shall be remote handling installation compatible
Target maximum annual dose 10mSv
Notes: Access will be controlled
Removal of hot spots
Installation of tempory shielding
Time limited
Increasingly over time
Forward area particularly unpleasant
Remote handling obligatory

## Slide 11: @ monolith face ready for remote handling
“Make space”
Sequence
Remove roof
Install platform
Install guide frame
Un fasten components with long tools
Cask (if required)
Remove
Notes: Remove highly active components
Activity of monolith

## Slide 12: Remote handling ready
Don’t reinvent the wheel !
Support structures
Fixations
Alignment 3 point
Guide rods
Service connections
Developed in collaboration with
Chopper are ready
Neutron guides are next !
Notes: A lot of thinking has already gone into this area
Internally
Also at other centers and more broadly in industry
We are not interested in re-inventing the wheel
Neither technical
Nor operationaly
Studied what has been done

## Slide 13
Logistics
(external to Bunker)
Hall Crane
Bunker Crane
Hand over zone
Bunker crane
20t winch
2t winch (requesting)
Coverage Acceptable
Hall Crane
30t winch
5t winch
Coverage Acceptable
Issues
Roof complexity = Many lifts
Storage within the hall is limited
Multiple handovers increase time required.

## Slide 14: Logistics “in bunker logistics”
Personnel access
‘C’ crane
Transport systems
Circumferential cranes number 2
Type : mono rail
Capacity 500kg
+ radial crane on an beam
Notes: Making the bunker work
Maintenance activities
Repair
Exploring
On the floor Activites

## Slide 15: Summing up
Strategy headlines
High reliability
Long service intervals
Standard interfaces
All designed for vertical removal
‘Critical’ systems equipped for remote handling
Common tooling
Notes: Starts with having a plan
Push for reliability
Building in Maintainability , manage risk
Priority on critical systems
Operational Integration
Tooling
Schedules
Spares
Work with experienced parters
It not going to be perfect use the precious early years to Learn
don't reinvent the wheel SNS
ISIS
Ensure future compatibility
upgrade path

## Slide 16: Thanks you for your attention !
Many thanks for contribution from the Bunker group and the Neutron Chopper Group

## Slide 17: Vibration mitigation
Reduce source terms
  Balance
  Magnetic bearings
Avoid coupling
  Rigid structures
  Rigid fixation
Isolate sources
  Support – Base
  Separate base plates (PPSc)
Attenuate
  Bolted & grouted base plate
Bolted & grouted baseplate
Spring loaded fixation
Reservation for Thin isolator
Rigid support
Notes: don't reinvent the wheel SNS
ISIS
Ensure future compatibility
upgrade path

## Slide 18: Target Safety related instrument failure
Current strategy “a layered and graded approach.”
Beam port a ‘thick’ window (Rated for normal safety case)
BP shutter a ’thick-ish’ window (v50)
Upstream side of high speed choppers a ’complimentary’ window
@ 11.5m a ’thick-ish’ window ‘fire break’
Metallic guides only upstream of ‘fire break’
BEAM PORT
GUIDE
GUIDE
Common shielding
SHUTTER
2
1
3
4
Safety related equipment
Normal equipment
Fire break

## Slide 19: Which equipment should be equipped to be remote handled ?
Equipment which represents a significant hazard to personnel during interventions (due to activation).
PPS choppers
Shutters & collimators
Some supports (?)
Equipment which requires regular intervention.
Overhaul or inspect period < 10 years.
Removed frequently i.e. is installed between R6 – R11.5
Equipment presenting a significant risk of breakdown. Reliability < 98%

## Slide 20: Which operation should be carried out using remote handling ?
Operations
Disconnection
  Beamline components, hold down and Services
Removal
  Fitment of cask if required
(Re)insertion
Repositioning or alignment (if required)
Reconnection

## Slide 21: Crane route
Bunker crane
Hall crane
Exchange zone
Storage zone

## Slide 22: Roof removal

## Slide 23: Full Short

## Slide 24: Full Short (Best case)
Number of Lifts 42
Shielding mass 180
Lift duration 16hr
Activities
  Beam port activation
  Instrument installation

## Slide 25: Full short (worst case)
Number of Lifts 47
Shielding mass 200
Lift duration 19hr
Activities
  Beam port activation
  Instrument installation

## Slide 26: R6-R9 (Front end)

## Slide 27: R6-R9 (Front End)
Number of Lifts 93
Shielding mass 80t
Lift duration 8h
Activities
  Maintenance
  Repairs
  Choppers
  Optics

## Slide 28: R9-R11.5 (Short Rear End)

## Slide 29: R9-R11.5 Short rear end
Number of Lifts 17
Shielding mass 100t
Lift duration 7hr
Activities
  Maintenance
  Repair
  all components

## Slide 30: Full Long

## Slide 31: Full long
Number of Lifts 93
Shielding mass 500t
Lift duration 16h
Activities
  Instrument installation
  Full overhaul

## Slide 32: Interventions - type 1A - Scheduled, PERIODIC (?)

## Slide 33: Type 1A Once in an (instrument) life time (10yrs) aka bunker annual events !
These activities are expected to be infrequent (periodicity >10yrs) when considered on a beam line by beam line basis.
Note when each bunker is considered as a system of 8-12 beam lines interventions become annual.
Interventions
Beam port activation
Instrument installation
Beam port deactivation
Instrument removal
Installation of neighboring instrument

## Slide 34: Interventions -type 1(b) schedule Regular (every shut-down

## Slide 35: Type 1B ‘the regulars’
These activities will be regular occurrences
These events will predominate work load of tech support during operation
Systems must be in place to enable these activities to be carried out minimising exposure to personnel
– ie minimising hands on.

## Slide 36: Scheduled & Regular
Activity Incidence over 10yr period
Beam port systems inspection 40
Neutron chopper maintenance 100
Optics realignment 20
Instrument shutter maintenance 20
Other component maintenance 60
Removal neighboring & Hot 80
Total 320

## Slide 37: Interventions - Type II unscheduled (BREAKDOWNS)
Breakdown, Screw ups, and other stuff that just happens

## Slide 38: Sub system level breakdown data
Source: ISIS
Note: excluding - Sample environments, Computing & Motion control
Causes often requiring access within heavy shielding for repair
Causes not requiring access within heavy shielding for repair
Chart: 
 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 1 | 1
 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 1
 | 3 | 1 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 3
Beam interruptions
Cycles (Normalized to 35 day)
CAUSE OF STOPPAGE
BLUE Chopper system
RED Other component
GREEN At request of I.S.
~ 3 ‘significant’ interruption per 100 beam days
+3.5 uses of shutters to effect ‘running repairs’
Notes:   Calculated failure rates for existing instruments and equipment
  And from that
  Factoring in the increase comlexity of ESS instruments tried to come up with some estimates for our own instruments
  Failure of a specific instruments every 3 or 4 years
  Failure of ‘An’ instrument each 35day cycle

## Slide 39: Unscheduled (breakdowns & unforseen)
Activity Incidence over 10yr period
Beam-port systems breakdown 4
Neutron chopper failure 40
Optics breakdown 20
Instrument shutter breakdown 20
Other component breakdown 60
Removal of hot components (neighboring work) 30
Accident scenarios (H2) 2
NOTE : Indicative figures – work in progress

## Slide 40: Availability & Source Schedule
Lengthening short breaks 3 > 5 days enables them to be used for repair.
Results in increased overall availability
Chart: 
| option A /B | Option C/D | Option E
225 BD | 66.4 | 93.4 | 81
215 BD + 4 Short breaks | 79.8 | 93.4 | 89
Notes: As a general rule unscheduled shut downs for repair are restricted to GREEN zones
Access to components in other zones must await scheduled ‘short shutdown’ – to cope with more complex interventions.
Longer interventions will be dealt with in two steps, with removal of equipment during one S-SD, repair off line and replacement at the next S-SD.
In the case of interventions requiring extended SD then shutters can be held close after start up of the facility.

## Slide 41: DC-LR installation
1.4m
.2.1m

## Slide 42
Worst case
Disc chopper Large rotor
Odin, Freia, (4-6 units)
N
2100
300
Assuming a rotor of 1800mm
The height of the unit is 2100mm
Assuming a lifting tool of 300mm
Gap to floor of 200mm
Free height require > 2600mm
200 gap

## Slide 43: Hook height
~1800 (700mm disc)
300mm (20t hook)
~2300
N BEAM
Gap 200mm

## Slide 44
Components and their placement
(with respect to the instrument in the beam line allocation document)
Components | Total number | In bunker | In bunker and serviced by monolith crane | Serviced by monolith crane but not in bunker
Over/Under Chopper | 23 | 16 | 16 | 0
Single Chopper | 33 | 19 | 16 | 2
Face-to-Face chopper | 26 | 20 | 19 | 0
Large rotor choppers (1-1.4m) | 7 | 7 | 5 | 0
Large rotor choppers (>1.4m) | 3 | 2 | 1 | 0
PPS | 6 | 4 | 4 | 0
Guide module | 218* | 201 | 165 | 10
Other | 12 | 3 | 2 | 2
Total | 391 | 272 | 228 | 14
*Every instrument is considered to have a single guide module from m30 onwards

## Slide 45: Installed systems What
Insert schematic of bunker components here

## Slide 46: Installed systems Neutron choppers
High speed double choppers 8-10 assemblies
PSS = 5 assemblies
High or Hot equipment

## Slide 47: DCSR (2PA-1-CRD) Conclusions
A hook height of 2.5m is a minimum for current equipment
2.7m prefered
2t – 5t lifting equipment would be strongly preferred.
Special mountings will be required to ensure installation compatibility