## Slide 1
ESHAC 5
Close out
04 October 2016

## Slide 2
How do you look upon the organisation of ESH and Q and the included cost estimate for the future operational phase?
The Committee believes that the proposed resources for operational radiation protection are adequate, except for the function of RPE where a single position is considered not to be sufficient and except for electronical maintenance of radiation monitors, where an extra 2 FTE could be necessary.
The Committee reiterates its continuous recommendation to have the necessary competence and resources for radiation physics calculations (shielding, activation ,...) within the EH&S Division. This should also allow maintaining a close link between the people doing the calculations and the people doing the operational radiation protection, which from experience in other large accelerator facilities is of key importance.
The Committee recommends that within the OHS Unit, one specialist in biology safety should be hired (can be one of the three positions showed for experiments safety.
The Committee recommends that the OHS Unit has specific competence in laser safety, in view of the expected large number of lasers, including class 3B and 4, on the instruments
The Committee recommends that a position for electrical safety is included in the OHS Unit. This position should in particular be in charge of electrical safety related to electrical user equipment.
The Committee recommends that the OHS Unit should include sufficient resources to manage the safety follow up of contractors, works accident analyses, ... From experience the Committee estimates the need for this as 3 (technician) FTEs.

## Slide 3
How do you look upon the organisation of ESH and Q and the included cost estimate for the future operational phase?
The Committee recommends that the OHS Unit should include sufficient resources for safety training. The Committee believes that the 1 FTE proposed by ESS would not be sufficient. This task could be shared with the 3 FTE mentioned in the previous recommendation. The Committee takes it for granted that the necessary IT resources are available within the ESS organisation to provide the necessary support for the administrative management of safety training databases (and IT needs to manage other ESH & Q issues).
The Committee recommends that within the ESS organisation chart, the required positions are foreseen to have “Experimental Hall Operators”, such as to have at least one person present 24h/7d. These people will have to assume a number of safety responsibilities, which clearly need to be defined. In particular, these people will be involved in the emergency response (injuries, fire alarms, ...).

## Slide 4
What is your opinion about the status of our SSM application work and the associated time schedule?
The Committee acknowledges the important amount of work accomplished by ESS since the submission of application 2 related to providing complementary information request by SSM.
However, the Committee is worried by the amount of complementary information still to be submitted. In view of the time needed by SSM to do the final review before delivering the permit for installation, all the complementary information should be submitted to SSM before the end of 2016, otherwise the general project planning may be delayed. The Committee believes this is a very tight schedule.
The Committee is also worried about the tight schedule to obtain the permit 3 without causing further general project delays. In view of this tight schedule the Committee believes that it is necessary that permit 3 should be split in two consecutive permits, the first permit for beam at the end of the 2 GeV linac but without beam on target, the second permit for beam on target.
The Committee believes that work on application 3 should start as soon as possible (i.e. basically today) to minimise the time between the delivery of the permit for the commissioning of the warm part of the linac (definition of basic requirmements from SSM) and the submission of the application for 2 GeV beam.
ESS should use lessons learned from the writing up of the application 2, in particular by nominating from the beginning a technical editor and by anticipating more time to write up the self assessment document.

## Slide 5
How do you look upon the work done for sustainability?
As the subject was not dealt with during the meeting, the Committee will examine the document provided by ESS and will reply in the final report.

## Slide 6
How do you look upon the work done for ODH?
The Committee commends the organization ESS has put in place with respect to ODH, in particular the creation of the ODH working group and the Cryogenic Safety Committee, with representations from the different divisions concerned. The Committee also commends the elaboration of the draft version of the ODH safety process and implementation document.
The definition of an escape lane via a hatch might need special justification in the safety file, as it is not good practice to escape into an area of higher radiation dose rate.
A technical review has been carried out with external experts leading to an optimized ODH detection implementation configuration. As a result it appears that the number of ODH monitors could be reduced by increasing the distance between them. The Committee wonders if this reduction has an impact on the criteria for triggering the evacuation alert and whether the false alarm aspect has been taken into consideration?
Please note that the cryogenic expert within the ESHAC was not present during the present meeting.

## Slide 7
Your comments regarding the schedule and needed work for the commissioning, especially the first part of the accelerator?
ESS needs to establish an operations organization (even if one person). Early in a project, Operations helps to define procedures, training, emergency response, commissioning plans, alarm management…Also can be trained in EPICS to create operations and commissioning screens, apps, …
The necessary posts in the organisation chart have to be filled, all necessary equipment, such as radiation monitors must be available, …
Commissioning should be on the agenda of next ESHAC.

## Slide 8
ESS responded to the recommendations from the last ESHAC meeting. The committee is invited to evaluate the responses and comment where appropriate.
(11) The Committee believes that all radiological risk studies (shielding and activation) should be supervised more effectively throughout the entire assessment process by the ES&H Division, implying the necessary expertise in radiation protection. ES&H Division should take the overall responsibility for the results, their documentation and verification of their implementation. This recommendation is directly related to the Committee’s worries about the delays in the hiring of the Radiation Safety Officer.
ESS notes that this has been a constant recommendation from the committee since its start of 2014. It has been discussed at the top management level of ESS but up to this date, the ESH organisation is not “responsible” for the calculation of the necessary shielding. The ESH Division is one of several reviewers that all shielding documents have to pass through. One has to bear in mind that for a neutron scattering facility, shielding calculation is not only a safety function but also has a functional meaning of “isolate” the instruments. This is a difference compared with a Synchrotron facility where shielding and radiation protection is more linked.
Comments from ESHAC: the committee notes the shielding co-ordination set-up as retailed in response to recommendation 11, which it still would prefer being integrated to the ESH services in terms of a radiation safety function.
The Committee reiterates its continuous recommendation to have the necessary competence and resources for radiation shielding calculations within the EH&S Division. This should also allow maintaining a close link between the people doing the calculations and the people doing the operational radiation protection, which from experience in other large accelerator facilities is of key importance.
(38) ESS shall follow up the lessons learnt at CERN with the adopted approach.
The ESH and Q division have to admit that the follow-up of the CE marking issue has not been in focus for the last year. Resistance from several stakeholders has led to a standstill in this question. With our new Ass Dir. in place, the issue will be raised again and we will follow-up the approach made by CERN more carefully.
Comments from ESHAC: the committee re-iterates its conviction that an adequate conformity assessment process, notably the availability of the hazard and risk analysis enabling system-wide compilation, modelled on European common practice (CE) should be in place for all project deliverables, whether in-kind or not.

## Slide 9
ESS responded to the recommendations from the last ESHAC meeting. The committee is invited to evaluate the responses and comment where appropriate.
(38) ESS shall follow up the lessons learnt at CERN with the adopted approach.
The ESH and Q division have to admit that the follow-up of the CE marking issue has not been in focus for the last year. Resistance from several stakeholders has led to a standstill in this question. With our new Ass Dir. in place, the issue will be raised again and we will follow-up the approach made by CERN more carefully.
Comments from ESHAC: the committee re-iterates its conviction that an adequate conformity assessment process, notably the availability of the hazard and risk analysis enabling system-wide compilation, modelled on European common practice (CE) should be in place for all project deliverables, whether in-kind or not.

## Slide 10
ESS responded to the recommendations from the last ESHAC meeting. The committee is invited to evaluate the responses and comment where appropriate.
(52) Incorporate a Controlled Access mode into the PSS proposal to SSM.
Currently the PSS system includes a controlled access mode, and this will be presented to SSM in November 2016. There is a planned meeting where ESS introduces the PSS strategy, analysis, software, hardware and shows the test stand that has been developed here at ESS.
Comments from ESHAC: the Committee hopes that SSM will be able to accept ESS’s proposal concerning the controlled access mode. The Committee is convinced that the incorporation of a controlled access mode will contribute to a reduction of the individual and collective radiation doses, as it will limit the number of tunnel searches.

## Slide 11
Top 6 recommendations
Radiation Monitoring System
The Committee recommends that ESS deals with radiation monitoring issues with the highest priority. Monitors must be specified and tested if no reference exists (in particular electronics most probably need to be able to cope with pulsed beams). The Committee was surprised to learn that in-kind contributions were again considered for radiation monitoring.
Radiation physics calculations (shielding, activation, ...)
The Committee reiterates its continuous recommendation to have the necessary competence and resources for radiation shielding calculations within the EH&S Division.
ES&H + Q resources for normal operations
The Committee recommends that ESS reviews its resource planning for ESH+Q taking into account the detailed recommendations formulated by ESHAC.
Licensing
See reply to question no. 2 of ESS charge.
CE certification / In kind contribution management
Installation
Commissioning plan
Resources + organisational and technical aspects
Shielding simulations for warm part of linac
Safety organisation for co-activity (with SKANSKA)