## Slide 1: Cold tests of bulk Nb cavities at CERN
Katarzyna Turaj
on behalf of
S. Barriere, A. Castilla, B. Frere-Bouniol, P. Fernandez Lopez,
K. Hernandez Chahin, A. Macpherson, K. M. Schirm,
N. Shipman, N. Stapley, A. Zwozniak, M. Wartak

## Slide 2: Bulk Nb Cavity Testing at CERN
1980
2017
RF cavity testing at CERN’s SM18 Facility:
  400 MHz Crab Cavities (Bare & Dressed) & Cavity Tuners
  704 MHz High Gradient Program,
  1.3 GHz cavity R&D

## Slide 3: CERN’s Bulk Niobium Program: Overview
• HL-LHC Crab Cavity
    RF validation of bare and dressed 400MHz cavities for SPS
    Cryomodule assembly and validation
    Installation into CERN SPS for validation with proton beams (2018)
• High Gradient Cavity Program
    Prepare and validate 5-cell ß=1 704 MHz cavities
    Bring 4 cavities to specification for assembly
    Assembly into a cryomodule for RF validation in CERN’s SM18 facility
• Process and Diagnostics development
    Cleanroom techniques and cavity assembly procedures
  RF surface preparation improvement
  Diagnostics for Quench localization,
  Bulk and surface RRR measurements
  Cooldown techniques, control of flux trapping

## Slide 4: Recent Activities: Key Points
Over the last 12 months
  Upgrade of infrastructure interleaved with cavity testing
  Ongoing upgrade of our 4m & 2m cryostats
  ~ 9 months for RF Testing
    Dominated by RF tests on Crab Cavities
HG Cavity Performance: Specification achieved
  Validated upgraded electro-polishing cathode design
      Chemistry now proceeding on remaining cavities
Cleanroom assembly process
  Process and infrastructure reviewed for Crab project
      Cavity handling and assembly more controlled & cleaner

## Slide 5: Schedule from first half of 2017
Last 12 months dominated by HL-LHC Crab activity
      High Gradient RF tests “squeezed in”
CRAB
High Gradient
Vertical Cryostat
Chemistry/ Assembly
Q1 2017
Q2 2017

## Slide 6: 2016 Results: High Gradient
SLHiPP-6
Improvements from cavity preparation significantly reduced field emission resulting in improved RF cavity performance

## Slide 7: 2017 Results: High Gradient
Continued improvement in preparation & test procedure => reduced field emission & reach specification
                Q0>1010 at Eacc=25 MV/m

## Slide 8: Performance Measurements
All scans performed in CW operation @ 2K
  Scan increments in watts => detailed scan at high power
Field emission above 20MV/m
  => Cavity assembly still to be improved
Moving to Self Excited Loop system
  Improvement of measurement uncertainties
Raw scan
Processed scan

## Slide 9: Quench studies
32 sensor array of OSTs
  Readout based on CEA-Saclay design
  Resolved issues with signal loss
    Signal lines in cryostat replaced (coax→ twisted pair)
  Cleaned up signal to noise ratio by isolating OSTs from cavity and cryostat infrastructure
Test new Transition Edge Sensors for quenches
  Offers improved positioning resolution
Hernán Furci’s talk

## Slide 10: Quenches - 3 Varieties Observed
Local Quench
  Slow quench
    Cell 1 equator
  2nd sound velocity
    V = 12.4 m/s
Global Quench
  Fast quench
    cell 3 equator
  No localisation
  2nd sound velocity
    V = 18 m/s
Double quench
  Small quench without cavity trip

## Slide 11: In-situ bulk RRR measurement
Standard 4-wire resistance measurement
  Measurement at 300K and at 10K
  Resistance at 10 K: R~0.2 µΩ
  Fixation of wires to cavity is critical
Procedure validated on Crab Cavities
Cavity | Resistance at 300K µΩ | Resistance @ 10K µΩ | RRR
DQW SPS 1 | 64.247 +- 0.282 | 0.189 +- 0.025 | 340 ± 45
DQW SPS 2 | 49.022 ±0.112 | 0.130 ± 0.005 | 377 ± 15
RRR on treated Nb samples = 290 ± 30

## Slide 12: Upgraded Cryo Insert - Validated in 2016
Improved insolation & cryostat now leak free
  Infux heat flow reduced & electrical grounding scheme
Cavity positioned lower in cryostat
  Larger head of liquid Helium above cavity
OLD INSERT
NEW INSERT

## Slide 13: Infrastructure: Cryostat insert upgrade
New ‘top’ of vertical cryostat commissioned
  Insert assembly, installation & load balancing
Vacuum pumping system replaced
  Increased diagnostics (including RGAs added)
  Pneumatic valves on pumping line:
    Avoids ambient B-field issues when opening shielding

## Slide 14: Improvements - Mobile Coupler
Switch to mobile coupler in Q3 2017
  Test power antenna re-designed for beam port
  Coupler designed to work in 2K liquid helium
  50 mm range of movement
    => can have critical coupling from 10K to 1.8K
  Will allow improved conditioning and cavity processing
Mass 7 Kg
335 mm
152 mm
Courtesy of E. Montesinos and team

## Slide 15: Cleanroom Practices
Post 2016 HL-LHC Crab Cavity cleanroom tooling review
  Re-tooling and improvement of ISO4 cleanroom infrastructure and practices
  Reinforcement of cleanroom floor and purchase of cleanroom lifter matched to project needs
  Removal of silver coated nuts/bolts from cavity assembly
      Using silicon bronze nuts for flange assembly
  Improved pre-assembly cleaning of pieces

## Slide 16: Other Improvements
Cleaner connection of cavity to insert pumping line
Reduction of contamination during 650ºC vacuum bake
    Use niobium caps on all ports during bakeout
Improved prep of cavity immediately prior to HPR
    Ultrasonic bath after final chemistry essential

## Slide 17: Outlook
HL-LHC Crab Cavity Program
  Vertical Cryostat RF tests continue through 2017
High Gradient Program: 4 Cavities
  HG1: Chemistry re done → Preparing for RF Test
  HG2: Achieved Specification: Now our reference cavity
  HG3: Bulk EP done. Sent to vacuum oven for 650ºC bake
  HG4: Waiting for EP
High Gradient Program: Cavity Testing
  RF tests to be interleaved with Crab program in 2017
  Combine with commissioning of mobile coupler
Bulk Niobium program
  Target: 1 RF cold test per month per cryostat
  2nd cryostat coming on line in Q3 2017

## Slide 18: Summary
CERN’s Bulk Niobium program
  dominated by Crab cavity activities at present
  … but still progressing with the High Gradient Program
  Soon to have 2 vertical cryostats dedicated to RF testing of bulk Niobium cavities
High Gradient: One cavity has reached specification
  Other three cavities to under go same treatment/testing
  Expect to finish initial test on all 4 cavities by Q1 of 2018
Ongoing developments
  Cataloging of quench behaviors and quench spot localisations prior to optical inspection
  Validated deployment of Self Excited Loop LLRF system and semi automation of test procedures
  Integration of mobile coupler into conditioning process

## Slide 19: Thank you