## Slide 1: WP11: CDS-SL CDR Instrumentation and safety devices
April 19th 2017

## Slide 2: Instrumentation - Overview
The CDS-SL valve box
TS circuit
Pressure sensor
Temperature sensors (x2)
Cryogenic control valve (x3)
Warm control valve
Cold mass circuit
Pressure sensor
Temperature sensors (x11)
Cryogenic control valve (x6)
Warm control valve

## Slide 3: Instrumentation - Overview
The CDS-SL end box
Ts circuit
Pressure sensors (x2)
Temperature sensors (x4)
Cryogenic control valve (x3)
Auxiliary lines
Pressure sensor
Warm control valve
Cold mass circuit
Pressure sensors (x2)
Temperature sensors (x16)
Level sensors (X2)
Cryogenic control valve (x2)
Electrical heaters (x6)
Insulating vacuum
Vacuum gauges (x2)

## Slide 4: Instrumentation - Summary
Sensors
Actuators

## Slide 5: Temperature sensors 1/2
Temperatures below 30K
Cernox™ 1050 - Lakeshore Cryotronics
Thermo-resistive sensor (Zirconium-oxynitride)
≈75Ω@300K; ≈10kΩ@2K
Good sensivity at cryogenic temperatures
Gold plated copper bobin
4 wires
Calibration @ IPNO
Brass screw M2.5
Copper support for temperature sensor
Surfaces which are greased by use of ApiezonTM type L or M? (1 MGy)
Temperature sensor
Cernox CX1050CU package
Face with the series number
CU package – Teflon free model
Stycast 2850FT + Catalyst 9 OK for ESS?
(cf. WP5 CDR)

## Slide 6: Temperature sensors 2/2
Temperatures above 30K and electric heaters
PT100 class A
Sensivity of about 0.4 Ω/°C [50K;300K]
≈108Ω@300K; ≈9.4 Ω @50K
Canister
Ts
Braze
Ts
Tc
Tw
Tw
Rth
Rth
From Lakeshore Cryotronics documentation
Surface which are greased with ApiezonTM type L or M? (1 MGy)
Temperature sensor
PT100
Ø=2.8mm
L=15mm

## Slide 7: Electrical heaters
Heating cartridges
Geometry | Diameter (mm) | 25.4+0-0.2
 | Overall length (mm) | 150+0-1% or-5mm
 | Heating length (mm) | 120
Heat transfer | Thermal load (W/cm²) | 1
Electricity | Supply voltage (V) | 48
 | Power (W) | 100
Leads | AXON (provided by IPNO) | 
Derived from a standard model (Vulcanic)
Temperature is monitored by means of two PT100 including redundancies

## Slide 8: Thermal anchoring
Valve box
End box
Thermal shield case
Cold mass case
Sensor
Multi channels
feedthroughs
Thermal anchoring
Sensor
Sensor thermal anchoring
TS thermal anchoring
Multi channels
feedthroughs
Thermal shield case
Cold mass case
Sensor
Multi channels
feedthroughs
Thermal anchoring
Sensor
Sensor thermal anchoring
Multi channels
feedthroughs
Electrical heater
Multi channels
feedthroughs

## Slide 9: Electrical connection
Internal instrumentation
EH: twisted pair
TT: two twisted pairs both twisted together
Four wires cable (copper)
Four channel electrical connector x2
This approach allows for :
Easy replacement of a component (sensor, prolongation, feedthrough)
An independent mounting of each part of the temperature sensor cabling
Sensor or actuator
Multi channels
feedthrough
Part 2: Cables prolongations
(~2 m)
connectors
Part3: Feedthroughs internal
Cables (~0,30 m)
connectors
Part1:
Sensor or actuator cable
(~0,75 - 1 m)
For reliability:
Electrical tests at room temperature (electrical continuity and insulation and sensor and cable resistance measurement)
Five thermal cycling with liquid nitrogen (cool-down shock and warm up)
Electrical tests again at room temperature

## Slide 10: Electrical connection 2/2
Internal cables
  Insulating material
  Conductor material
Multi-channels feedthrough and internal connectors
Radiation resistance > 5,0.104 Gy
TT (from vacuum vessel to TS) – Manganin, Phosphor/brass, Copper – AWG TBD (supplier)
EH – Copper – AWG TBD (supplier)
From AXON radiation resistant cables for scientific and nuclear applications documentation
Example of a multi-channels feedthrough (LEMO HGP.5S.324.CLAP(E)) mounted on a KF DN50 flange.
Thermometer internal connector (LEMO FGG.0S.304.ZLAT) (left)
Electrical heater internal connector (LEMO FGG.1S.302.ZLAT) (right)

## Slide 11: Liquid helium level sensors
Superconducting level gauge
Superconductive
Nb-Ti filament
Heater
Supply current 70mA
≈4.55 Ohm/cm (normal state)
Expected active length : 60cm
Electrical connections
Multi channels
feedthrough
End box insulating vacuum
End box phase separator
Two twisted pairs both twisted together
Vacuum vessel wall

## Slide 12: Pressure sensors 1/3
Requirements
  Radiation dose: 5,0.104Gy remote electronic sensor
  Tightness of the cryogenic circuits fully welded construction
Technical solution
Manufacturers and models
  Baumer PDAS (used and qualified at CERN)
  HBM P3
  Honeywell TJE and STJE (suggested by ESS)
Measurement range
  Thermal shield circuit: 0-25 bara
  Cold mass circuit and SV relief line: 0-6 bara
Max measure P > MAWP
 allows for pressure monitoring if a critical event occurs
  Measurement principle based on strain gauges

## Slide 13: Pressure sensors 2/3
Main characteristics

## Slide 14: Pressure sensors 3/3
Cost estimation (IPNO (WP4+WP11) + CEA (WP5) need)
  « PDAS sensors are no longer manufactured »
  After discussions with Baumer Special manufacture for ≈ 100 units is possible…
  …but expensive

## Slide 15: Control valves positioner
Radiation compliant solution : SIEMENS SIPART PS2
  Positioner for pneumatic control valves with remote electronics
  Industrial solution available on the market
Position sensor
(potentiometer)
Inlet piezoresistive control valve
Outlet piezoresistive control valve
Compressed air supply
To the atmosphere
Qualified at CERN
On top of each valve
Drawer: 5, 10, 15 valves
WPXX: Cryogenic C&C
WP11: CDS-SL
Positioner hosts :
Position sensor
Piezoresistive valves (x2)
Drawer hosts:
Electronics of the control loops for 5, 10 or 15 valves
Communication module (Profibus PA)

## Slide 16: Vacuum components 1/2
The vacuum instrumentation port (x1)
  On the top of the vacuum vessel of the End Box of the CDS-SL
  KF DN40 (ISO 2861) made in stainless steel
  Equipped with a manual vacuum valve
  One primary (Pirani) and one secondary
  (Cold cathode) gauge
EB control port
EB instrumentation port

## Slide 17: Vacuum components 2/2
The vacuum control ports (x14)
  On the top of the vacuum vessel of each VB and the EB
  KF DN40 (ISO 2861) made in stainless steel
  At least 4 ports equipped with a manual vacuum valve
  Other ports equipped with KF DN40 (ISO 2861) blank-off flange and tightened with an EPDM o-ring and a collar
VB control port

## Slide 18: Vacuum gauges
UHV Vacuum gauges for Accelerator Physics from MKS (suggested by ESS)
Mechanical interface : CF40
CF to KF adaptors needed

## Slide 19: Vacuum valves
Angle or in-line manual vacuum valves main properties (TBC)
Sealing material: EPDM

## Slide 20: Safety devices - Overview
The CDS-SL end box
TS circuit
Main supply line (SV93)
Main return line (SV94)
He circuit
Main supply line (SV91)
Main very low pressure line (SV92, RD92)

## Slide 21: Safety devices - Overview
The CDS-SL valve box
TS circuit
Branch return line (SV60)
He circuit
Branch supply line (SV01)

## Slide 22: Safety devices - Summary
Safety device | Location | Device type | Supplier | Supplier description | Set pressure | Req min. d0 | Size and d0 | PN | Maximum opening hysteresis or uncertainty | MAWP
 |  |  |  |  | (barg) | (mm) | (mm) |  |  | 
SV91 | He Supply (End Box) | Safety valve, spring-loaded | LESER (TBC) | 4414.467x xx | 4 -> 3.75 | 35 | DN50 - d0=46 PCTFE | 40 | +/- 10% TBC | 5
RD92 | VLP (End Box) | Reverse Buckling Rupture Disk (welded) | Witzenmann (TBC) | BDRI.XX100.4,5 (TBC) | 3.7 -> 4.5 | 45 | 45 < d0 < 100 | - | +/- 5% | 5
SV92 | VLP (End Box) | Safety valve, spring-loaded | LESER (TBC) | 4594.217x xx | 3 -> 3.8 | 14 | DN20 do=13 PCTFE | 40 | +/- 10% TBC | 5
SV93 | TS supply (End Box) | Safety valve, spring-loaded | LESER (TBC) | 4594.217x xx | 21 | 9 | DN20 do=13 PCTFE | 40 | +/- 10% TBC | 24
SV94 | TS return (End Box) | Safety valve, spring-loaded | LESER (TBC) | TBD4414.464x xx | 21 | 20 | DN25 do=23 PCTFE | 40 | +/- 10% TBC | 24
SV01 | He branch supply (Valve Boxes) | Safety valve, spring-loaded | Herose (TBC) | 06850 (TBC) | 3,75 | 6 | 10 | 6 | +/- 10% | 5
SV60 | TS branch return (Valve Boxes) | Safety valve, spring-loaded | Herose Circle Seal (TBC) | 06850 (TBC) | 21 | 6 | 10 | 25 | +/- 10% | 24
List of the CDS-SL safety devices
ESS change
WUST/ESS sizing
Not critical: small quantity of fluid in a pipe + small surfaces

## Slide 23: Safety devices – Branch lines
Po (bar) | Pa (bar) | ρV (kg/m3) | Surf. écran (m²) | Heat load density (W/cm²) | Heat load (kW) | Mass flow (kg/s) | k hélium | Crit, Pressure ratio | Pa/P0 | Ψ | A (mm2) | α | Dmin (mm)
22 | 1,1 | 17,39 | 11 | 1500 | 16 500 | 0,05 | 1,67 | 0,487 | 0,050 | 0,51 | 18 | 0,73 | 4,79
Thermal shield safety valve sizing SV60:
  Considered worst scenario: loss of insulation vacuum
    air leak into the vacuum vessel of the CM ( = 50 mm)
    air fully condensates onto the CM thermal shield surface covered with 30 MLI
    considered heat power density : 1.5 kW/m² [Lehmann, 1978]
    for 1 (CM + jumper) thermal shield (~11 m²): 19.5 kW
  Helium properties considered (supercritical helium):
  ISO 21013:
    Mass flow rate:
    Sonic flow considered as
    Min. relief cross section:
Req. d0  6 mm
P (bara) | T (K) | (J/kg) |  (Pa.s) |  (kg/m3)
22 | 58 | 3.18105 | 7.3310-6 | 17.39
P (bara) | T (K) |  |  (Pa.s) |  (kg/m3)
22 | 58 | 3.18105 | 7.3310-6 | 17.39

## Slide 24: Safety devices – Branch lines
Burst disk:
  For the prototypes (WP4 and WP5 + test Spoke valve box):
    CF flange
    Possibility to have a welded interface (Witzenmann)
    To be updated (also depends on the interface with the vent line)
Safety valves:
  Interface to be confirmed but:
    LESER: thread + external weld possible
    HEROSE: welded pipe possible
  with lever arm for control actions if welded

## Slide 25: Safety devices
Pressure scale:
  Presented during the safety review at Lund (2016/06/09th)
  Same principle but here updated
18.0: minimum pressure to allow the closing of the relief valve (0.9 x 20.5  18,0 bar)
21.0: set pressure of the safety valve (Pset  5%)
24,25: maximum pressure for a 100% opening of the relief valve
(1,1x 22.0  0.55 bar)
 20.0 (0.95 x Pset)
 22,0 (1.05 x Pset)
Gauge pressure
(barg)
Absolute pressure
(bara)
PS = MAWP = 24
Pressure operation in the temperature range 35-55 K
19.5
19
25.25
Max. operating pressure
Relief pressure
(SF relief line)
 [1 atm ; 1.1 bar]
Opening area of the relief valves
 ongoing: checking the tolerances of the safety valves with the suppliers: +/- 10% (total) TBC

## Slide 26: Thank you for your attention
CSP12
28 Novembre 2014

## Slide 27: Electrical connection
Internal cables
  Insulating material
Radiation resistance > 5,0.104 Gy
From AXON radiation resistant cables for scientific and nuclear applications documentation

## Slide 28: Control valves positioner 1/2
Radiation compliant solution : SIEMENS SIPART PS2
  Industrial solution available on the market

## Slide 29: Control valves positioner 2/2
Radiation compliant solution : SIEMENS SIPART PS2
  Industrial solution available on the market

## Slide 30: Instrumentation - Overview
Thermal shield circuit and auxiliary lines
  Valve box
  End box
TS circuit
Pressure sensor
Temperature sensors (x2)
Cryogenic control valve (x3)
Warm control valve
Ts circuit
Pressure sensors (x2)
Temperature sensors (x4)
Cryogenic control valve (x3)
Auxiliary lines
Pressure sensor
Warm control valve

## Slide 31: Instrumentation - Overview
He circuit and insulating vacuum
  Valve box
  End box
Cold mass circuit
Pressure sensor
Temperature sensors (x11)
Cryogenic control valve (x6)
Warm control valve
Cold mass circuit
Pressure sensors (x2)
Temperature sensors (x16)
Level sensors (X2)
Cryogenic control valve (x2)
Electrical heaters (x6)
Insulating vacuum
Vacuum gauges (x2)