## Slide 1: PDR of Master Oscillator
Anders Svensson 2017-04-28

## Slide 2: Outline
Introduction
Concept and interfaces
Requirements
Reliability (Redundancy)
Rack layout and location
Design and implementation
Prototype Results
Benchmark
Project Plan

## Slide 3: MO - Introduction
Primary timing and frequency source for ESS
Accurate timing information to the ESS timing system (EVG) based on 88.0525 MHz clock, 1 PPS and UTC information from GPS
Highly stable RF signals to LLRF and BPM, Syncronized to GPS
Located in the Gallery next to EVG and distributes reference to Phase Reference Line
High availability

## Slide 4: Concept
GPS disciplined Rubidium source
Crystal Oscillator (OCXO) with low close in phase noise
DRO running at 704 MHz for best phase noise performance
Distribution Unit (DU) delivers clock and RF signals
EPICS interface for monitoring and control

## Slide 5: External interface
Interface | Type | Level | Details
GPS antenna input | 50 Ω coaxial | See x.xx | TNC jack, L1:1575.42 MHz
GPS Ref output | 50 Ω, BNC female | +13 dBm ± 2 dB | 10 MHz sinusoidal
AC supply | AC, 50 Hz | 230VAC | AC with UPS
RF signals 704.42 MHz 352.21 MHz | 50 Ω, N female 50 Ω, N female | +10 dBm +10 dBm | sinusoidal sinusoidal
Timing Generator 1 PPS 88 MHz UTC | 50 Ω, BNC female 50 Ω, N female IPv4/IPv6, RJ-45 | Positive TTL pulse + 10 dBm TBD - | 20 us, 1 ms, 100 ms or 500 ms widths sinusoidal or square Time and date information
Ethernet | IPv4/IPv6, RJ-45 | - | For monitoring and control

## Slide 6: MO – PRL interface
Interface requirements to PRL:

## Slide 7: MO performance requirements
Trade off between cost, complexity and performance
SNS performance comparison
System impact: LLRF and BPM
LO generation
PA noise

## Slide 8: SNS performance
Phase noise performance:
Source: THE SPALLATION NEUTRON SOURCE RF REFERENCE SYSTEM, M. Piller, M. Champion, M. Crofford, H. Ma, ORNL/SNS, Oak Ridge, TN 37831, USA
L. Doolittle, LBNL, Berkeley, CA 94720, USA
Phase drift:
Reference Line:
Long term phase drift:
Part of MO phase drift

## Slide 9: MO impact of LLRF performance
  MO reference jitter degrades LLRF stability performance
  Reference also used by LO generation units
LO to Cavity
1 MHz
REF to Cavity

## Slide 10: PRL Amplifier impact
MO output: +10 dBm (10mW)
PRL-line input power: ~+53 dBm (200W)
Amplifier adds spurious and noise, R&S BC500
Input
Output: 56 dBm

## Slide 11: ESS performance
Jitter RMS:
Offset [Hz] | Phase Noise [dBc/Hz] |  | 
 | 88 MHz | 352 MHz | 704 MHz
10 | -110 | -98 | -92
100 | -123 | -111 | -105
1k | -148 | -137 | -135
10k | -156 | -150 | -155
100k | -162 | -156 | -160
1M | -162 | -158 | -165
Target values:
BW | 88 MHz | 352 MHz | 704 MHz
10 Hz.. 1 MHz | 42 fs | 37 fs | 35 fs
100 Hz.. 1 MHz | 26 fs | 17 fs | 13 fs
DRO performance
MO phase drift target:
between 352 and 704 MHz: < ± 0.05° (normalized to 704 MHz)

## Slide 12: SNS performance (ESS)
Phase noise performance:
Source: THE SPALLATION NEUTRON SOURCE RF REFERENCE SYSTEM, M. Piller, M. Champion, M. Crofford, H. Ma, ORNL/SNS, Oak Ridge, TN 37831, USA
L. Doolittle, LBNL, Berkeley, CA 94720, USA
Phase drift:
Reference Line:
Long term phase drift:
Part of MO phase drift
Offset [Hz] | Phase Noise [dBc/Hz] |  | 
 | 88 MHz | 352 MHz | 704 MHz
10 | -110 | -98 | -92
100 | -123 | -111 | -105
1k | -148 | -137 | -135
10k | -156 | -150 | -155
100k | -162 | -156 | -160
1M | -162 | -158 | -165

## Slide 13: Reliability
Redundancy -> high availability
Preventive maintenance
Design considerations
Analysis including amplifiers (PRL)

## Slide 14: Rack location and layout
3 racks located in MBL-070 to get balanced power distribution of PRL
Two identical racks for reliability reasons, and third for EVG and switch panel
Rack1
Rack2
Rack3

## Slide 15: GPS
GPS disciplined Rb source
Ethernet and 1 PPS
Antenna outside the Gallery building and RF coax to receiver input

## Slide 16: OCXO PLL
Generates 100 MHz from 10 MHz reference provided by Rb source
OCXO PLL design

## Slide 17: DRO PLL
Generates 704.42 MHz from 100 MHz reference provided by OCXO
DR in cavity
Stabilized temperature

## Slide 18: Distribution Unit (DU)
Provides 352 MHz and 88 MHz signals
Ethernet interface to EPICS using XT-Pico module
RF levels and temperature monitoring
15 IO’s TTL
LEDs
Voltage monitor
I2C bus
5V, 1A
Serial number on chip
Connector
Divider AD9515
XT-Pico
Distribution Unit

## Slide 19: OCXO and DRO PCB
OCXO
10 MHz
704 MHz
DC input
Monitoring
Cavity feeds
µ-controller
DRO PLL
Voltage regulators
Same PCB for simplicity
  size 155 x 155 mm
  4-layers

## Slide 20: Prototype results
OCXO PLL
DRO PLL

## Slide 21: Benchmark - overview
Frequency generation provided by Wenzel
  SNS MO provided by same vendor
  3U standard 19” rack mount
  Monitoring over EPICS not supported

## Slide 22: Benchmark – performance and cost
Frequency | 88 MHz | 352 MHz | 704 MHz
Offset [Hz] | Phase noise [dBc/Hz] | Phase noise [dBc/Hz] | Phase noise [dBc/Hz]
100 | -125 | -121 | -115
1k | -153 | -149 | -143
10k | -165 | -166 | -160
100k | -165 | -171 | -165
HW cost | Wenzel | ESS design
3 units | $134.000 | €20.000
Performance sligthly better than ESS design
HW cost
GPS and switch panel monitoring not included

## Slide 23: Project Plan - Design
PDR: 28-April
CDR: 27-Sep

## Slide 24: Project Plan - Procurement
Procurement: starts after CDR and ends 10-Jan
SAT: end of January 2018