## Slide 1: Status of coating development
Dr. Cyrille Thomas
www.europeanspallationsource.se
23 October, 2017

## Slide 2: Outline
Large Collaboration
Objectives and Requirements
Down-selection
Results
Baseline and First Materials
Schedule Coating development

## Slide 3: Large Collaboration
ESS BI: project lead, planning, measurement and characterisation, X-ray, neutron, photo-luminescence, vacuum, thermal
ESS Target Division and NSS: Material scientists: advise, environment and material lifetime calculations, R&D luminescent materials
Oslo University: support with measurements and characterisation, XRD, low energy proton beam, photo-luminescence, thermal
HV: Develop coating process, prepare samples for research, establish industrialisation of the coating
BNL: BLIP, high energy proton Irradiation test (under RADIATE collaboration)
GSI: heavy ion facility, collaboration with BI physicists, share expertise on luminescence materials
SNS: studies on a high power target, share expertise, common interest of finding a new luminescent material
Stony Brook University: coat SNS target, research and collaboration on SNS coating and new coating materials
J-Parc: proposed collaboration, low energy ions facility,
Los Alamos National Lab: Post Irradiation Examination: XRD, High energy Proton beam
5/8 w irradiation -> 1y delay
Possible in the Jap. Fisc. Y18
DTU: Irradiation plasma coated Al2O3:Cr

## Slide 4: Objectives and Requirements
Objectives:
  Deliver a luminescent material as the source for the imaging systems
  Deliver an industrialised coating process (control, reliability, quality assurance, manage schedule)
Requirements:
  Luminescence performance (yield, spectrum, lifetime) under environment constraints (radiation, temperature)
  Lifetime (material properties) larger or equals Coated Component’s lifetime
  Controlled coating process (quality assurance, schedule)

## Slide 5: Objectives and Requirements
Luminescent material characteristics satisfying the requirements:
The photon yield: predict signal intensity, signal to noise ratio and related image quality
The emission spectrum: the photons wavelength (or energy) is of prime importance.
  It should be in an efficient range for the readout image systems
  It should be on a different from the proton beam induced fluorescence of the (residual) gas environment.
The luminescence lifetime: should be less or much less than 1µs in order to perform a still image of the beam on the moving target
Radiation tolerance: the luminescence properties should remain sufficient to permit minimal imaging performance along the life of the coated component
The temperature dependence: the luminescence yield at should be understood. The SNS coating has a strong dependence on temperature, and at 473K its luminescent yield is dramatically reduced
Luminescence from secondary particle: neutron, gamma, and other hadron particles may distort the beam density profile
Mechanical properties: compatibility with the coating process and strength of the coating
Vacuum compatibility: outgassing rate of the material should be acceptable for the vacuum system

## Slide 6: Down-selection
The down selection process will go through several phases:
Phase 1: Material selection and evaluation of process compatibility:
  Material identification, compatibility with thermal spray, compatibility with mechanical requirements.
Phase 2: Refinement of the deposition process, performance of factory floor testing:
  Sprayed sample and refinement of the coating process, and verification of the photo-luminescence
Phase 3: Photoluminescence Characterisation:
  Photoluminescence study: time resolved and spectra, vacuum compatibility, photoluminescence temperature dependence.
  As complementary studies, and for later irradiation test, the most promising material candidates should also have structural measurements records. This can be done by means of X-ray Diffraction techniques, Neutron Diffraction Techniques, and eventually SEM imaging.
Phase 4: Irradiation with no or minimal activation:
  Proton luminescence yield and spectrum, low energy proton beam, but also at high energy beam.
  Characterise radiation tolerance: thermo-mechanical studies, and luminescence properties
Phase 5: Irradiation with significant activation:
  Irradiation with neutron and protons, aim to mimic radiation environment of ESS
  Qualify the material for high power target after PIE, phases 3 and 4
  Phase 6: Industrialization:
  Final phase: documents the industrialised process for the preparation and coating of the selected material on the TW, the PBW, and the Dump (Dump surface and / or screens)

## Slide 7: Down-selection
 |  |  | Requirements:
Phase 1 | Material selection and evaluation of process compatibility | HV: materials coated on Al, steel, photo-luminescent checked | 3
Phase 2 | Refinement of the deposition process, performance of factory floor testing | ESS and Oslo Univ. conduct further characterisation of the produced samples: 1st validation of the process | 1, 3
Phase 3 | Photoluminescence Characterisation | ESS and Oslo Univ. : Complete characterisation, XRD, photon-luminescence (spectrum and lifetime), proton-luminescence | 1
Phase 4 | Irradiation with no or minimal activation | Oslo Univ. , GSI, J-Parc, possibly COSY and LANSCE: Low and High energy proton beam, heavy ion beam | 1, 2
Phase 5 | Irradiation with significant activation | BNL, SNS, DTU proton irradiation at high dose: qualify the material for high power target environment Synchrotron (Diamond, NSLSII), PNNL Post Irradiation Examination (PIE): structure and composition of the irradiated material, photo and proton –luminescence, SNS, LANL other particles luminescence | 2, 1
Luminescence properties
Coating lifetime
Industrial process

## Slide 8: Results
Sprayed Brodmann powder:
    Luminescent coating process established with plasma gun and combustion gun
Luminescence properties:
  Yield:
    photo yield measured
    proton luminescence spectra and images, direct yield not measured due to limited control system deployed in Olso
  Luminescence lifetime: measured, decay 90% in 5ms
  Temperature dependency:
    R12 lines intensity is 70% less at 200C
    Integrated spectrum (600nm – 800nm) 50% at 200C
  Vacuum compatibility checked
  Radiation tolerance: not yet done!
    BLIP irradiation 5/8 complete, delayed to 2018, PIE to follow
      Qualification for 7 different plasma coated Al2O3: Cr stuctures, and Y2O3
    DTU irradiation:
      Start with 4 samples: 2 showing best yield, one intermediate and worst yield (XRD correlated)
      Follow-up with first combustion coated samples, and lanthanides doped Y2WO6

## Slide 9: Photo-luminescence Yield
Plasma coated samples
Oslo Cyclotron: 80nA, 15ms, gain 0dB
LED 365nm, 12mW, 3mm2:
Luminescent radiant power: 50µW/Sr/mm2
Calibrated image and spectral power-intensity
Done under BI developed EPICS control
Power meter (NIST)
LED
image
Spectrum
Camera vendor demo software

## Slide 10: Luminescence lifetime
Test-Bench developed and tested
  HV plasma coated samples characterised
  Comparison with SNS coated samples
Sample | t (ms) | a | t (ms) | a | t (ms) | a | t (ms) | a
HV-7 | 0.05 | 0.68 | 0.38 | 0.21 | 3.4 | 0.052 | 20 | 0.061
HV-5 | 0.06 | 0.64 | 0.38 | 0.27 | 3.4 | 0.052 | 21 | 0.035
SNS | 0.13 | 0.73 | 0.86 | 0.04 | 0.9 | 0.14 | 9.3 | 0.09
Mod. Source
L2
D1
BS
S
F
D2
Pico8
L1
L3

## Slide 11: Temperature Dependency
In vacuum and in air measurements
  Vacuum compatible
  Narrow band: 75% decay @200 ◦C
  Broad band: ~50% decay @200 ◦C

## Slide 12: Radiation tolerance
BLIP:
  5/8 weeks of irradiation done
    Complete in 2018
  PIE: @PNNL, opening the cassette, XRD, TEM, SEM
    Synchrotron complementary measurements: localised XAF, XRD vs. dose/DPA
    Results by the end 2018
DTU:
  Mitigate the large delay from BLIP
  Access to the sample during irradiation:
    Luminescence characterisation as function of dose/DPA
    Complement data with synchrotron
  Results right after irradiation, and complement data within 3 months
  Qualification of new materials

## Slide 13: HiRadMat
Series of samples: Broadman qualified and new formula, Yttria, Zirconium stabilised Yttria
  Received pulses with similar beam current density to ESS
  Verify high energy luminescence
  Verify high mechanical strength of the coating

## Slide 14: New Luminescent Materials
Yttrium base luminescent materials:
  High luminescent yield (literature)
  Used at high temperature (literature: Y2WO6)
  Potential for reduction and developing of extra spectral lines (O lines), used for imaging when main lines have decayed
Y2WO6 doped with lanthanides:
  Prepared at ESS, luminescence characterised (Yield)
  Next steps are to spray and irradiate

## Slide 15
Initial results with doped Y2WO6:
  Photo-luminescence
  XDR
  Proton luminescence
Y2WO6 : | Er | Eu | Nd | La | Pr | Sm
Power (nW) | 982 | 799 | 692 | 898 | 913 | 981
Yield | 1,8% | 1,4% | 1,2% | 1,6% | 1,6% | 1,8%
Oslo Cyclotron

## Slide 16: Baseline and First Materials
Baseline coating for the TW, PBW, DumpL:
  Brodmann Al2O3:Cr new formula
    Combustion / plasma gun coating on PBW, DumpL
    Combustion gun coating on TW
Already qualified at SNS, but it will be fully qualified following a conservative approach
Next step is in the procurement phase, contracting for coating TW, PBW and DumpL screens
Research for new materials continues:
  Ahead development for the second coated elements
  Done using the 6 month replacement of PBW, and DumpL screens

## Slide 17: Schedule Coating development
Mock-up Ready
2020
2021
2019
2018
2017
BLIP Irradiation
Uncomplete
DTU
PBW Spray
Review: Coating Process Approved
BLIP: cassette ready
PIE (LANL)
Beam cancelled!
TW (BEW) Spray
Dump screens Spray
Coating Industrial process
CDR
BLIP Irradiation
Complete
DTU
New coatings
SAR 1