## Slide 1: Beam on Target Diagnostics
Dr Cyrille Thomas
Beam Diagnostics Engineer
www.europeanspallationsource.se
18 November, 2017

## Slide 2: Outline
Concept and Overview
Beam tuning and monitoring
  Case of the Dump
  Case of the Target

## Slide 3: Concept and Overview
Suite of instruments to measure beam properties on target:
    Size, position (IMG)
    Current density distribution (IMG, GRID)
    Beam outside defined aperture (APTM)
  Scale time: Intra-pulse, pulse, cumulative pulses
Objectives:
  Support beam on target tuning:
    establish nominal operation
    protection
  Support beam on target operation (production):
    monitor
    protection

## Slide 4: Beam on Dump
BPM: phase, position, trajectory
  Blind to long bunch
BCM: peak current along pulse
APTM: charge loss into the defined aperture
  Measure all charges from MHz sample to minutes
  Can be biased by secondary electrons
IMG: beam transverse 2D profile, peak current density distribution
  Detects small beam size
  Limitation in dynamic range: too large beam size may not be seen

## Slide 5: Beam on Dump
Tuning
for every pulse:
Verify beam size predicted by the model
MPS on: check beam size larger than sth (mode dependant)
Check beam current on aperture (protection)
Check trajectory and divergence (protection)
APTM
Screen 1
Screen 2
Screen 3
Screen 4
X1, Y1, sx1, sy1
DX, DY, Dsx, Dsy
X2, Y2, sx2, sy2
Dump
Image 1
Image 2
IMG
APTM
Current / sectors (MS /s
MS/s
Temperature / Thermocouples
1µA
0.8µA
3µA
4µA
350k
360k
310k
314k

## Slide 6: Beam on Target
Accelerator region
Target region
Heine Thomsen (ISA, Aarhus University)
TW
PBW
BEW
PBIP
GRID
APTM
IMG (x2): current density
CO
A2T and Raster Magnets
BPM
BCM
NPM
APTM: beam aperture (protection)
BPM: phase, position
BCM: peak current
NPM: beam profile, size, positon
WS: (low power) beam profile, size
APTM

## Slide 7: Beam on Target
Beam parameters:
Power: 2MW
Peak Current: 62.5mA
Avg Current: 2.6mA
Pulse: 2.86 ms
PBW
BEW
H. GRID
V. GRID
FFT peaks
V: 29kHz
H: 39kHz
APTM
1µA
0.8µA
3µA
4µA
350k
360k
310k
314k
1µA
0.9µA
2µA
4µA
350k
360k
310k
314k
PBW
PBIP
IMG
Max peak density: 50%
Max peak density: 30%

## Slide 8
Thank you