## Slide 1
K. Nakajima
Neutron Science Section
MLF, J-PARC Center
III. Instrument technology and experiences
Performance of instruments (intensity & resolution in comparison to design, some of background issues)
The 2nd ESS–J-PARC Collaboration Workshop
18-19 January, 2018
ESS, Lund, Sweden

## Slide 2: Neutron Instruments at MLF
Nano-Structure Group
・Reflectometer
BL16: SOFIA Horizontal
BL17: SHARAKU Polarized
・Small&Wide-ANS
BL15: TAIKAN
Pulsed Neutron Application Group
・Neutron Source Diagnostic & Test Port
BL10: NOBORU
・Neutron-Nucleus Reaction Meas. Inst.
BL04: ANNRI
・Fundamental Physics
BL05: NOP
・Neutron Imaging
BL22: RADEN
Operation
Operation
Operation
Operation
Operation
Diffraction Group
・Single Crystal Diffractometer
BL03: iBIX for Bio-Materials
BL18: SENJU
・Powder Diffractometer
BL08: S-HRPD Hi-Resolution
BL09: SPICA for Battery Mat.
BL20: iMATERIA Hi-Intensity
BL11: PLANET Hi-Pressure
・Total Scattering Diffractometer
BL21: NOVA
・Engineering Diffractometer
BL19: TAKUMI
Operation
Operation
Operation
Operation
Operation
Operation
Operation
Operation
2008~
2011~
2008~
2011~
2008~
2010~
2008~
2008~
2008~
2011~
2010~
2008~
2008~
Spectroscopy Group
・Chopper Spectrometer
BL01: 4SEASONS Hi-Intensity
BL12: HRC Hi-Resolution
BL14: AMATERAS Cold-Neutron
BL23: POLANO Polarized
・Inverted Geometry Spectrometer
BL02: DNA
・Neutron Spin-Echo
BL06: VIN_ROSE
Operation
Operation
Operation
Operation
2008~
2008~
2009~
2011~
Commissioning
Construction
Operation
2008~
Operation
2015~
Operation
2017~
MIZE
NRSE

## Slide 3
Neutron Instruments at MLF
－ Operation 19.5
BL22: RADEN
BL23: POLANO
BL06:
VIN ROSE
－ Comissioning 0.5
Empty lines 2 (BL07, 13)
－ Construction 1
Notes: 装置について、
BL18の立ち上げが進み、微小試料構造解析を実証したこと
BL03は検出器数が14台（去年）から３０台になったこと
昨年のNACではBL08とBL19のissue がspeed up to restartでしたので、
これらが復旧して、研究を開始したことがトピックスです。
BL09はコミッショニングを開始、リートベルト解析までができるように
なったこと（米村さんが中島さんに送ったファイルをCCします）
BL11はコミッショニングが終了、科研費メンバーにより、実験が開始
されたこと、来年度共用装置に移行すること（このファイルは持って
いると思います）
BL21は外部資金の期間が終了したこと
サイエンスについて、
BL03はトランスサイレチン以降ですが、課題審査部会報告にありますが、
詳しく説明を受けないと、使えません。
BL08は目ぼしい成果でないと思うかもしれませんが、重要です。
英語にしましたので、添付します。
BL19 成果あり（LPSO結果：日本語しかないので、英文を依頼中）
BL20 成果あり（英文を依頼中）
BL21去年はPhysical Review Lett.の結果を出していない思います。
（英文を依頼中）

## Slide 4: Neutron Instruments at MLF
HRC
BL12
4SEASONS
BL01
↓
AMATERAS
BL14
↓
Experimental Hall No. 1
Experimental
Hall No. 2
DNA
BL02
↓
POLANO
BL23
VIN ROSE
BL06
S-HRPD
BL08
SPICA
BL09
TAIKAN
BL15
SOFIA
BL16
iMATERIA
BL20
NOVA
BL21
SHARAKU
BL17
RADEN
BL22
PLANET
BL11
NOBORU
BL10
NOP
BL05
ANNRI
BL04
iBIX
BL03
TAKUMI
BL19
SENJU
BL18

## Slide 5
I. Performance of Instruments in Comparison to Design
(PICD)
- Neutron Flux
- Resolution

## Slide 6: PICD: Neutron Flux
The performance of the neutron source
F. Maekawa et al., NIMA 620, 159 (2010)
Scattering from Mica.
The performance of the neutron source examined at BL10-NOBORU
  Intensity : well agreed with neutronics calculation result within 20% error.
  Profile : well agreed with neutronics calculation result, except small tails at long wave length region.

## Slide 7: PICD: Neutron Flux
Examples of well-agreed cases
BL12 HRC guide section L~5 m SN
S. Itoh et al., NIMA 631, 90 (2011)
BL11 PLANET guide section L~12 m M&SN
T. Hattori et al., NIMA 780, 55 (2015)
BL03: roughly agreed
BL15: obs./calc. = 96%
BL18: obs./calc. = 84%
BL20: obs./calc. ~ 70%
BL21: roughly agreed
BL06 VIN ROSE (NRSE) guide section L=20m KUR
Other examples
M. Hino et al., J. Nuc. Sci. Tec., 54, 1223 (2017)

## Slide 8: PICD: Neutron Flux
K. Nakajima et al., EPJ Web of Conferences 83, 03011 (2015)
BL14 AMATERAS guide section L~26m I
observed / simulation > 10%
Examples of not-so-well-agreed cases
R. Kajimoto et al., Phys. Soc. Jpn. 80, SB025 (2011)
BL01 4SEASONS guide section L~13m I
BL19 TAKUMI guide section L~30 m M
observed / simulation = 65%
observed / simulation = 60%
BL02: Obs./Calc. = 61%
BL06(MIZE): Obs./Calc. = 65%
BL16: Obs./Calc. ~ 10%*
Other examples

## Slide 9: Beam-Transport of AMATERAS
Sample
(=focal point)
Moderator
27.7m
2.3m
90mm
4m
focal point
Straight Section
18.2m
Elliptical
Section
Transit
Section
8m
Vertical: Straight + Elliptical
Moderator
Sample
30mm
Horizontal: Curved Guide (R=2Km)
Curved Section from 7.2m to 26.8m
Combination of 3Qc&3.8Qc super mirror
Sample
Moderator
Bottom & Top 3.0Qc
Both Sides 3.0Qc
Shutter
Section
Biological Shielding
2.3m
4.3m
7.0m
Bottom & Top 3.0Qc
Both Sides 3.8Qc
Bottom & Top 3.0Qc
In-Side of Curve 3.0Qc
Out-Side of Curve 3.8Qc
Curved Section
18.2m
Curved Section
26.8m
Straight
Conversing
(Elliptical)
Bottom & Top 3.8Qc
In-Side of Curve 3.0Qc
Out-Side of Curve 3.8Qc
Bottom & Top 3.8Qc
Both Sides 3.8Qc
J-PARC
(Japan)
3.8Qc
Xenox
(France)
3&3.8Qc
Mirrotron
(Hungary)
3Qc
Laser Tracker
Corner Cube Refrector

## Slide 10
Sources of the beam loss
Intensity Problem at AMATERAS
Broken-alignment (movement after alignment)
At TAKUMI, intensity was partly recovered by re-alignment.
AMATERAS employed the same manner & the same company.
Borofloat-glass mirror problem
At the construction of DNA, it was found that borofloat-glass mirror reflectivity was 60% of expectation.
AMATERAS employed borofloat-glass mirror in 10 m of total guide section.
Others
Beam-transport re-work on AMATERAS (2012-2015）
re-alignment
exchange of mirror
re-alignment: all sections after the biological shielding. (including refurbishment of equipment)
Exchange of borofloat-glass mirror.

## Slide 11: Re-Work on AMATERAS BT
AMATERAS beam-transport re-work (2014)
K. Nakajima et al., proc. of ICANS-XXI (JAEA-Conf-2015-002, 306 (2016))

## Slide 12: PICD: Resolution
AMATERAS
(
)
ch
ch
ch1
mod
3
1
ch1
mod
ch
mPS
m
→
2
1
1
t
t
L
L
L
L
t
t
t
D
+
D
-
-
+
D
=
D
D
-
-
Monochro. Chop. = 300Hz, P-Shap. Chop. = Open, P-Rem. Chop. = Open
Monochro. Chop. = 300Hz, P-Shap. Chop. = Open, P-Rem. Chop. = 300Hz
Monochro. Chop. = 300Hz, P-Shap. Chop. = 300Hz, P-Rem. Chop. = 150Hz
Pulse Removal
Chop.
Energy resolution at elastic position well agrees with analytic calculation.
K. Nakajima et al., J. Phys. Soc. Jpn. 80, SB028(2011).

## Slide 13: PICD: Resolution
4SEASONS
Energy resolution and Q resolution
well agreed with calculation.
Calculation of resolution in single-
crystal experiment at choppers.
  Simulation + Analytic calculation
K. Iida et al., JPS Conf. Proc. 1, 014016 (2014)
R. Kajimoto et al., proc. of ICANS-XXI (JAEA-Conf-2015-002, 319 (2016))
R. Kajimoto et al., arXiv:1705.04869 (AIP Conf. Proc., to be published )
C4H2I2S
Al2O3

## Slide 14
Ei = 42 meV | Observed | Calculated | Simulation
0 meV | 4.72 meV | 3.8 meV | 4.59 meV
14 meV | 2.97 meV | 2.3 meV | 2.79 meV
30.5 meV | 1.49 meV | 1.0 meV | 1.14 meV
(HH0)
Ei = 42.05 meV
(HHL) L=1.5 r.l.u.
Qa / r. l. u. = 1.42 - 1.58
Qb / r. l. u. = 1.40 - 1.60 or 2.40 – 2.60
Qc / r. l. u. = -0.08 - 0.08
Ex. CH01&CH03 open； CH02 both 150 Hz; CL04&CL05 1st frame
Sample： Single-Crystal of Cu
Energy Resolution of AMATERAS

## Slide 15
II. Performance of Instruments Background
- AMATERAS
- 4SEASONS
4SEASONS 2008~
AMATERAS 2009~
R. Kajimoto et al., Phys. Soc. Jpn. 80, SB025 (2011)
K. Nakajima et al., Phys. Soc. Jpn. 80, SB028 (2011)

## Slide 16: Improvement of S/N
2009.6
~20 kW
First measurement
2009.11
~100 kW
Addition of several shielding materials
Addition of detectors around the beam center
2012.01
~100 kW
Installation of the T0 chopper
Further addition of shielding materials
Magnetic Excitations in CuGeO3 at 4SEASONS

## Slide 17
Background problem at 4SEASONS
R. Kajimoto et. al., JPS Conference Proceedings 8, 034013 (2015)
Background Caused by Continuously Emitted High-Energy Neutrons
This type of background becomes serious on 4SEASONS.
chopper open
chopper at p/4 position
direct beam position
High-energy neutrons continuously emitted.
High-energy neutrons penetrate B4C beam catcher and 10B disk-chopper blades.
Even in the case of a Hg-target neutron source.
Possible sources of elements: 9Li, 17Ni, 8He, 16C...

## Slide 18: Counter background at AMATERAS
At the beam-line
a) Jacket of (curved) guide mirror (20t SUS)
b) Beam line shielding (concrete & iron )
c) Chopper housing (> 50t SUS)
d) Polyethylene

## Slide 19: Counter background at AMATERAS
Around the scattering chamber
a) Cd inside the tank
b) B4C mortar
c) B4C & H3BO3 floring
d) B4C liner inside the get-loss tube

## Slide 20
Background at AMATERAS
Is background at AMATERAS really low?
― when accelerators off
Sources of background observed here...
Electrical noise
Neutrons from the floor
→Neutrons moderated in the floor concrete. Origin is cosmic ray.
Not only those!!!

## Slide 21: Background at AMATERAS
3. Powder diff. of B4C of narrower
4. Air scattering (~1Pa)
5. Other γ-ray
(from chopper disk?)
2. γ-ray from optical devices
6. Unknown spot-like scattering
7. T0 Burst
1. Electrical noise, cosmic ray

## Slide 22
beam-off
Background problems at AMATERAS
Time-independent background
~30 Counts h-1PSD m-1
It doesn't depend of whether beam is on/off.
Background caused by prompt γ-ray.
Cd and, in sometime, a sample itself cause prompt γ-ray, which still can be detected by 3He counters.
We are trying to replace Cd on the direct beam line with other martials.
K. Nakajima et. al., proc. of ICANS-XXI; proc. of ICANS-XXII (to be appeared in JPCS )

## Slide 23
Example of γ-ray Background
S. Ito et al. Nat. Com. 8, 235 (2017)
Adjust LLD
From users' paper
Cd at DW Narrower
Cd at UP Narrower

## Slide 24
Summary
Performance in Comparison to Design
Intensity:
Majority of instruments achieve > 80% of expectation.
Some others suffered from BT problems.
Resolution:
In the case of chopper spectrometers, resolution well can be explained by analytic calculation.
Background at 4SEASONS & AMATERAS
Continuous efforts have been payed.
High-energy neutrons can be detected @ 4SEASONS.
Several sources of backgrounds are observed @ AMATERAS.

## Slide 25
List of Publications
General overview of MLF instruments
K. Nakajima et al., Quantum Beam Science 1, 9 (2017)
Instrumental papers shown here
T. Hattori et al., NIMA 780, 55 (2015) PLANET(BL11)
S. Itoh et al., NIMA 631, 90 (2011) HRC(BL12)
M. Hino et al., J. Nuc. Sci. Tec., 54, 1223 (2017) VAN ROSE(BL06)
R. Kajimoto et al., Phys. Soc. Jpn. 80, SB025 (2011) 4SEASONS(BL01)
K. Nakajima et al., Phys. Soc. Jpn. 80, SB028 (2011) AMATERAS(BL14)
F. Maekawa et al., NIMA 620, 159 (2010)
Performance of the MLF neutron source
AMATERAS beam transport background issues
K. Nakajima et al., proc. of ICANS-XXI (JAEA-Conf-2015-002, 306 (2016))
K. Nakajima et. al., proc. of ICANS-XXII (to be appeared in JPCS)
4SEASONS background issues
R. Kajimoto et. al., JPS Conference Proceedings 8, 034013 (2015)
Beam transport design
K. Nakajima et al., EPJ Web of Conferences 83, 03011 (2015)
4SEASONS resolution issues
K. Iida et al., JPS Conf. Proc. 1, 014016 (2014)
R. Kajimoto et al., proc. of ICANS-XXI (JAEA-Conf-2015-002, 319 (2016))
R. Kajimoto et al., arXiv:1705.04869