## Slide 1: FLUX TRAPPING STUDIES ON LOW BETA CAVITIES
D. Longuevergne, P. Duchesne
IPN Orsay
SLHIPP-8, Uppsala, 12-13 June 2018

## Slide 2: OUTLINE
FACILITIES AND EQUIPMENT @ IPNO
  The cryostat and magnetic configurations
  The 3 types of cavities
MAGNETIC SENSITIVITY MEASUREMENTS
  On SPIRAL2 Quarter-Wave Resonator
  On MYRRHA Spoke Resonnator
  On ESS Spoke Resonnator
ANALYSIS
  Cooling conditions
  Geometrical effects
  Field dependence
  Temperature dependence

## Slide 3: The cryostat and magnetic configurations
Magnetic shielding
  Passive : 0.6 mm mu-metal all around the cryostat : B < 25 mG
  Active : 3 coils inside mu-metal sheets
The vertical cryostat
Residual field simulation with no compensation coils
P. Duchesne

## Slide 4: The cryostat and magnetic configurations
Magnetic configurations
OPTIMAL :
Bv < 10 mG
Bh < 10 mG
40 mG :
Bv ~ 40 mG
Bh < 10 mG
75 mG :
Bv ~ 75 mG
Bh < 10 mG
110 mG :
Bv ~ 110 mG
Bh < 10 mG

## Slide 5: The 3 types of cavities
QWR for SPIRAL2
SPIRAL2 QWR
F0 = 88 MHz
Bpk/Eacc = 10.5 mT/MV/m
r/Q = 515 
G = 33 
MYRRHA Spoke
F0 = 352 MHz
Bpk/Eacc = 7.3 mT/MV/m
r/Q = 217 
G = 109 
ESS Spoke
F0 = 352 MHz
Bpk/Eacc = 6.9 mT/MV/m
r/Q = 426 
G = 130 

## Slide 6: OUTLINE
FACILITIES AND EQUIPMENT AT IPNO
  The cryostat and magnetic configurations
  The 3 types of cavities
MAGNETIC SENSITIVITY RESULTS
  On SPIRAL2 Quarter-Wave Resonator
  On MYRRHA Spoke Resonnator
  On ESS Spoke Resonnator
ANALYSIS
  Cooling conditions
  Geometrical effects
  Field dependence
  Temperature dependence

## Slide 7: Magnetic sensitivity
OBSERVATIONS :
  Very low sensitivity to vertical field (~ 0.006 n/mG)
  Sensitivity to horizontal field (~ 0.05 n/mG)
  Geometrical dependence of magnetic sensitivity
  Linear dependence of surface resistance coming from flux trapping
SPIRAL2 QWR
@ 88 MHz, @ 1.7K, Smag= 0.056 n/mG

## Slide 8: Magnetic sensitivity
OBSERVATIONS :
  Low sensitivity to vertical field ~ 0.06 n/mG @ 2K ~ 0.054 n/mG @ 4.2K!
  Linear dependence of surface resistance coming from flux trapping
MYRRHA Spoke
@ 352 MHz, Smag= 0.12 n/mG @ 2K
Smag= 0.16 n/mG @ 4.2K

## Slide 9: Magnetic sensitivity
OBSERVATIONS :
  Low sensitivity to horizontal field
  ~ 0.09 n/mG @ 2K
  ~ 0.075 n/mG @ 4.2K!
  Temperature dependence of magnetic sensitivity
  Linear dependence of surface resistance coming from flux trapping
ESS Spoke
@ 352 MHz, Smag= 0.12 n/mG @ 2K
Smag= 0.16 n/mG @ 4.2K
2K
3K
3.5K

## Slide 10: OUTLINE
FACILITIES AND EQUIPMENT AT IPNO
  The cryostat and magnetic configurations
  The 3 types of cavities
MAGNETIC SENSITIVITY RESULTS
  On SPIRAL2 Quarter-Wave Resonator
  On MYRRHA Spoke Resonnator
  On ESS Spoke Resonnator
ANALYSIS
  Cooling conditions
  Geometrical effects
  Field dependence
  Temperature dependence

## Slide 11: ANALYSIS
On MYRRHA Spoke :
  SLOW COOL-DOWN (T ~ 0.5K) FAST COOL-DOWN (T ~ 4K)
  With fast cool down, magnetic field difference before/after transition is 3 times bigger!
  Expect a significant change in surface resistance ?
Cooling conditions

## Slide 12: ANALYSIS
On MYRRHA Spoke :
  Between slow and fast cool-down, a difference of 0.2 n measured over 6 n.
  No flux expulsion is observed but field re-distribution after transition
Cooling conditions

## Slide 13: ANALYSIS
Geometrical effect
Flux trapping :
Only normal component to the surface is trapped (SP2 QWR is a good example).
How to predict magnetic sensitivity of a structure with simulation code :
  Need to know normal vector of each mesh => mapping of trapped flux
Sensitivity is measured experimentally from Qo measurement : power measurement
  Trapped flux induces dissipations only in RF magnetic field regions

## Slide 14: ANALYSIS
Geometrical effect
 | Trapped flux | RF surface currents | Sensitivity to magnetic field
SPIRAL2 QWR – vertical field |  |  | 
SPIRAL2 QWR – horizontal field |  |  | 

## Slide 15: ANALYSIS
Geometrical effect
n/mG |  | Transverse sensitivity | Beam axis sensitivity | Vertical sensitivity
SPIRAL2 QWR @4.2K @ 88 MHz | Calculated | 0.048 | 0.048 | 0.011
 | Measured | 0.05 |  | 0.006
 | Error (%) | 4 |  | -50
MYRRHA SPOKE @4.2K @352 MHz | Calculated | 0.088 | 0.076 | 0.069
 | Measured |  |  | 0.05
 | Error (%) |  |  | 20
ESS SPOKE @4.2K @352 MHz | Calculated | 0.078 | 0.066 | 0.084
 | Measured |  | 0.075 | 
 | Error (%) |  | 9.5 | 
Theoretical
0.08 n/mG
Theoretical
0.16 n/mG
Theoretical
0.16 n/mG

## Slide 16: ANALYSIS
Field dependence
R0 = Rres + RBCS
Rs = R0 + R1*Eacc

## Slide 17: ANALYSIS
Field dependence
Possible explanation : collective weak pinning
  Cornell’s analytical approximation (D. Liarte) :
  R1 decreases « linearly » with temperature
  R1 decreases « linearly » with external field
  R1 affected by heat treatment
D. Liarte, CORNELL

## Slide 18: ANALYSIS
Temperature dependence

## Slide 19: CONCLUSION
Cooling conditions
  Magnetic field variation during transition is measured
  The variation depends significantly on cooling speed (T gradient)
  But no significant change of magnetic sensitivity
Geometrical effect
  As only the normal component of magnetic field is trapped, the geometry should have a significant impact on sensitivity
  A good prediction of sensitivity is calculated by the model presented here
Field dependence
  The linear dependence of the surface resistance due to trapped flux looks correlated to Rbcs+Rres and not only Rres!
Temperature dependence
  The residual resistance caused by flux trapping is temperature dependent!

## Slide 20: THANKS FOR YOUR ATTENTION

## Slide 22: Quarter-Wave Resonnator
Designed at IPNO
1 spare cavity fabricated by RI
Helium tank in titanium
QWR for SPIRAL2
Material : Bulk Niobium
 = 0.12
F0 = 88 MHz
T : 4.2K
Eacc : 6.5 MV/m
Bpk/Eacc = 10.5 mT/MV/m
Epk/Eacc = 4.9
r/Q = 515 
G = 33 

## Slide 23: Spoke Resonator
Designed at IPNO
2 prototypes fabricated by E. Z. Zanon Spa
Received in december 2015
Helium tank in titanium
SSR for MYRRHA
Material : Bulk Niobium
 = 0.37
F0 = 352 MHz
T : 2K
Eacc : 6.4 MV/m
Bpk/Eacc = 7.3 mT/MV/m
Epk/Eacc = 4.3
r/Q = 217
G = 109

## Slide 24: Spoke Resonator
Designed at IPNO
2 prototypes fabricated by E. Z. Zanon Spa. 1fabricated by SDMS
Received early 2015
Helium tank in titanium
DSR for ESS
Material : Bulk Niobium
 = 0.5
F0 = 352 MHz
T : 2K
Eacc : 9 MV/m
Bpk/Eacc = 6.9 mT/MV/m
Epk/Eacc = 4.3
r/Q = 426
G = 130
Pick-up Port
Coupler port
Bellows
4 HPR Ports
E field
H field

## Slide 25: ANALYSIS
On Spiral2 QWR :
Cooling conditions